## Tabelle1
Nexcimer Exception ID (CVX-300 Fault Code) | Name | Error Popup Text | Subtext | Exception Type | Icon | Recovery Mechanism | Translation Rationale
101 | NO_ENERGY | No energy detected at sensor | Position catheter tip 1-2" (2.5-5.0cm) from detector surface | Alert | Caution_Triangle | N/A | No energy was detected at the calibration energy sensor but was evident at the PES within the first 10 pulses of a calibration sequence.
102 | VESSEL_SATURATED | Requested Fluence level unavailable | Lower the requested Fluence setting and try again. Contact Customer Service to schedule maintenance. | Warning | None | Lower desired fluence | The system is incapable of supplying enough energy to meet the specified fluence during PROCEDURE Mode.
103 | CATHETER_INPUT_SATURATED | Error 103: System Failure | Contact Customer Service if problem continues. | Error | Caution_Triangle | N/A | The energy required at coupling plane exceeds 180mJ.
104 | Removed with ECO040762 | N/A | N/A | N/A | N/A | N/A | N/A
105 | VES_OUT_OF_RANGE | Error 105: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The VES energy reading is 15% out of range.
106 | SENSOR_MISMATCH | Error 106: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The readings between the laser vessel and pulse energy sensors, normalized for optical path losses, do not match within a limit established as the sum of 2% VES variance, 2% PES variance, 2% attenuation variance and an additional 25% margin, adjusted for the expected attenuation.
107 | ENERGY_OUT_OF_RANGE | Error 107: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The estimated energy delivered to the patient (during PROCEDURE Mode) exceeds ±20% of the fluence setting for 100+ pulses.
108 | MAX_INVALID_PULSES | Error 108: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | More than 10 of the last 100 pulses were invalid in PROCEDURE MODE.
109 | PREVENTIVE_MAINTENANCE | Maintenance Reminder | Schedule maintenance for this device in the next 2 weeks. | Warning | Maintenance | N/A | The voltage required to generate the maximum fluence for the inserted catheter exceeded 975V measured during calibration. The system is still functional, but a gas exchange should be scheduled within two weeks.
Reminder will be displayed at the completion of calibration if condition is met.
110 | REMOTE_INTERLOCK_OPENED | Error 110: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | Close facility doors | The facility door was opened.
200 | BAD_CATH_INSERTION | Error 200: | Remove and reinsert catheter. Contact Customer Service if problem continues. | Error | Caution_Triangle | Remove Catheter | The catheter coupler switches were not activated in unison, and the fault can only be cleared by re-inserting the catheter.
201 | INVALID_CATHETER_ID | Error 201: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | Remove Catheter | The catheter inserted did not have an entry in the catheter default parameters table. Recovery occurs when the offending catheter is removed.
210 | SHUTTER_NOT_CLOSED | Error 210: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The safety shutter was commanded to the closed position, but the position indicator does not confirm it.
211 | SHUTTER_NOT_OPEN | Error 211: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The safety shutter was commanded to the open position, but the position indicator does not confirm it.
220 | ATTENUATOR_FAULT | Error 220: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The attenuator motor driver reported a fault.
221 | ATTEN_POS_FAULT | Error 221: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The attenuator position does not match the commanded value.
230 | HVPS_SAFETY | Error 230: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The HVPS safety shutoff pin or driver output is not at the expected level.
231 | HVPS_ERR | Error 231: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The HVPS reported a fault.
240 | BLOWER_FAULT | Error 240: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The blower motor controller indicated a fault.
250 | FOOTSWITCH_DISCONNECTED |  | Insert a footswitch plug into either connection point. | Error | Caution_Triangle | Connect Footswitch | The footswitch is not connected. The fault is cleared when the footswitch is connected.
Error 250:
Contact Customer Service if the problem continues.
251 | TOO_MANY_FOOTSWITCHS | Error 251: | Two footswitches detected. Remove one footswitch to continue. | Error | Caution_Triangle | Unplug one of the two footswitches | Two footswitches detected
Contact Customer Service if problem continues.
252 | FOOTSWITCH_INVALID | Error 252: System Failure |  | Error | Caution_Triangle | N/A | The footswitch is in an invalid state - Fault detected (either both NC/NO or neither).
Remove and re-connect footswitch.
Contact Customer Service if problem continues.
260 | CAN_RX_OVERFLOW | Error 260: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The CAN bus hardware buffer has overflowed and messages may have been lost.
270 | MISSING_PARAMS | Error 270: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The reporting module did not have the necessary parameters to begin firing.
300 | INVALID_STATE_ENERGY | Error 300: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The Energy Control Module reported the wrong Mode and/or State.
301 | INVALID_STATE_FIRING | Error 301: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The Firing Control Module reported the wrong Mode and/or State.
302 | INVALID_STATE_SAFETY | Error 302: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The Safety Control Module reported the wrong Mode and/or State.
303 | INVALID_STATE_UI | Error 303: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The UI Control Module reported the wrong Mode and/or State.
310 | MODE_STATE_CHANGE_FAULT | Error 310: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | One or more modules did not indicate readiness to complete the state or mode change after being commanded to change modes or states.
320 | NOT_IN_READY_STATE_CALIBRATE | Laser Not Ready | Release footswitch and press “Calibrate” to enable the laser. | Alert | Caution_Triangle | N/A | While in "Catheter inserted" screen. Cannot press the footswitch while system in in STANDBY State - release the footswitch and press CALIBRATE before pressing footswitch.
321 | NOT_IN_READY_STATE_PROCEDURE | Laser Not Ready | Release footswitch and press “Set to ready” to enable the laser. | Alert | Caution_Triangle | N/A | Cannot press the footswitch while system in in STANDBY State - release the footswitch and press READY before pressing footswitch.
322 | NOT_IN_READY_STATE_FAILED_CALIBRATE | Laser Not Ready | Release footswitch and press “Try again” to enable the laser. | Alert | Caution_Triangle | N/A | While in "Unsuccessful calibration" screen. Cannot press the footswitch while system in in STANDBY State - release the footswitch and press CALIBRATE before pressing footswitch.
400 | ENERGY_HB_TIMEOUT | Error 400: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The Session Control Module did not receive a heartbeat message from the Energy Control Module within 1.25s
401 | FIRING_HB_TIMEOUT | Error 401: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The Session Control Module did not receive a heartbeat message from the Firing Control Module within 1.25s.
402 | SAFETY_HB_TIMEOUT | Error 402: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The Session Control Module did not receive a heartbeat message from the Safety Control Module within 1.25s
403 | SYSTEM_HB_TIMEOUT | Error 403: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | A module did not receive the system heartbeat message from the Session Control Module within 1.25s.
404 | UI_HB_TIMEOUT | Error 404: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The Session Control Module did not receive a heartbeat message from the UI Control Module within 1.25s
500 | SELF_TEST_FAILED | Error 500: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The reporting module failed self-test.
No fault tolerance time is applicable in that POWER_UP_TIMEOUT would occur.
501 | POWER_UP_TIMEOUT | Error 501: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The system failed to initialize within 30 seconds.
502 | OPER_PARAMS_CORRUPT | Error 502: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The operational parameters were found to be corrupt or are not initialized.
503 | SW_VERSION_INCOMPATIBILITY | Error 503: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The SW Version from each of the modules does not match the list of versions associated with the System Version.
504 | FRAM_INIT_FAIL | Error 504: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The FRAM has failed to initialize properly. The Maintenance Application must send a message to explicitly reset the FRAM.
505 | FRAM_WRONG_MODE | Error 505: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | N/A | The system is in the wrong mode while performing a write operation.
600 | TEMP_HIGH _PCN_FANS | Approaching Temperature Error (600) | Allow the system to cool and try again. If problem continues contact Customer Service. | Alert | Caution_Triangle | Temperature Sensors drop below "High" range (system cools down) | The temperature sensors on PCN fan temperature switch. When the fans are off the system provides optical output signal. The signal is HIGH or LOW, active LOW.
Only issued once at transition from HIGH (INACTIVE) to LOW (ACTIVE).
601 | TEMP_HIGH_HVPS_OUTPUT | Approaching Temperature Error (601) | Limit pulse count, or allow the system to cool to avoid Error. | Alert | Caution_Triangle | Temperature Sensors drop below 45°C range (system cools down) | The temperature of the temperature sensor located on the PIM is above 50°C (measured with an accuracy of ± 1°C).
Only issued once at transition from <50°C to >50°C.
602 | TEMP_HIGH_E-TUB | Approaching Temperature Error (602) | Limit pulse count, or allow the system to cool to avoid Error. | Alert | Caution_Triangle | Temperature Sensor drops below 45°C (system cools down) | The temperature of the temperature sensor located on the XCM is above 50°C (measured with an accuracy of ± 1°C).
Only issued once at transition from <50°C to >50°C.
610 | VESSEL_OVER_TEMP | Error 610: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | Vessel Over-temp Switch deactivates (system cools down). Reset by Field Service | The laser vessel over-temp switch has been activated. The fault clears when acknowledged by the user or when the over-temp sensor deactivates.
611 | BLOWER_MOTOR_OVER-TEMP | Error 611: System Failure | Contact Customer Service for assistance. | Error | Caution_Triangle | Temperature Sensors drop below 75°C range (system cools down) | The temperature sensors on Blower Motor is above 100°C (measured with an accuracy of ± 1°C).
612 | TEMP_VERY_HIGH_HVPS_OUTPUT | Error 612: System Failure | Allow the system to cool and try again. Contact Customer Service if problem persists. | Error | Caution_Triangle | Temperature Sensors drop below 50°C (system cools down) | The temperature of the temperature sensor located on the PIM is above 58°C (measured with an accuracy of ± 1°C).
613 | TEMP_VERY_HIGH_E-TUB | Error 613: System Failure | Allow the system to cool and try again. Contact Customer Service if problem persists. | Error | Caution_Triangle | Temperature Sensors drop below 50°C (system cools down) | The temperature of one or more temperature sensors (Chassis, XCM) is above 58°C (measured with an accuracy of ± 1°C).