## Sheet1
ActionID | Problem Code | Root Cause | Fault Displayed | Reported Problem Level 1 | Reported Problem Level 2 | US Reporting Decision | Complaint Status | Valid Complaint? | State/Province | Country | Device Serial | Nature of Complaint | Field Service Summary | Patient/User Outcome | Field Service CRN Number | When was complaint detected? | Did the device malfunction? | Days_Open | Fault Code
161111 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Wiring | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100001 | 15 seconds after you switch the Philips Laser System on there is the message: "system error 110 - call the service" | FSE found a loose connection causing an error 110, repaired the connection and completed calibrations. System is operating properly. | No Patient Present, | 20210307-BV | Not during a case | Yes | 71
161115 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Shutter Assembly | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | GERMANY (DEU) | 100005 | Fault code 500. Issue is resoved over the phone and after rebooting the fault code does not show up anymore. We will schedule a courtesy visit. | FSE found the shutter did not open, replaced. Completed calibrations and system is operating properly. | No Patient Present, | 20210001 | Not during a case | Yes | 71 | 500
161702 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Shutter Assembly | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100006 | fault 500 | FSE could not duplicate however the logs indicate the safety shutter did not open, replaced. Completed calibration and system is operating properly. | No Patient Present, | 20210318-BV | Not during a case | Yes | 60 | 500
161915 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Operator Error | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100008 | We turned on the PLS, we performed the calibration with the sample sheath available and it was ok.After that the physician calibrated the PLS with the Glidelight without any problems.While we were touching the button “Ready” appeared ERROR 108 and we read this message: “ERROR 108: contact the Philips service” on the control panel. Then we removed the Glidelight’s connector, we re-inserted it and performed a second calibration, the PLS worked very well for the rest of the procedure.During the second procedure, we didn’t have any problem. We finished both the extraction procedures without any delay. | Non dispatch phone fix. Tech support determined the operator was selecting the ready button and depressing the footswitch simultaneously. System is operating properly. | No Injury, | 20210320-BV | During Preparation for Use | Yes | 57 | 108
163714 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Shutter Assembly | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | GERMANY (DEU) | 100005 | Fault code 500 | FSE was not able to duplicate the reported problem. Did note the shutter did not open in the logs, replaced shutter and safety control module. Completed calibrations and system is operating properly. | No Patient Present, | 20210365-BV | Routine Service/Not during a case | Yes | 47 | 500
163715 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Unable to replicate complaint | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | UNITED KINGDOM (GBR) | 100003 | Error code 106 was displayed | FSE was not able to duplicate the reported problem, the system calibrated and operated properly on arrival. | No Injury | 20210366-BV | During Preparation for Use | Yes | 47 | 106
166710 | The laser system does not provide reliable performance | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | N/A | ITALY (ITA) | 100008 | Dear all, I would like to inform all of you about the issue we are facing with the new Philips Laser System installed in Perugia in June. On Friday, for the 4th time, I’ve been informed that the physicians had issues with the calibration of the sterile probe: the system showed an error message of the failed calibration.They tried to disconnect and reconnect the probe and start again the calibration process. They also turned off the system but the same error was displayed.In order to complete the procedure they had to use a new probe and everything was ok. I’m supporting our dealer Newtech with the Customer Complaint forms in order to transfer all the necessary info to the quality department and to have the replacement devices. On that note, I think that a FSE visit could be useful to check from a technical point of view that the system is operational in conformity with technical specifications.As we know, this is a very important customer and I would like to prevent the potential loss of revenue due to the unreliability of the system. @Scarcella, Martino, are you already trained to perform this kind of intervention or should we consider the hypothesis to schedule Stephan for this activity? Any kind of suggestions are welcomed to fix this problem . Thank you all, Charlotte GuarrasiClinical Application Specialist Central ItalyImage Guided Therapy Devices Philips, Excelsiorlaan 41, 1930 Zaventem, BelgiumEmail charlotte.guarrasi@philips.com Mob +39 347 983 40 33www.philips.com/IGTdevices | FSE found optical rail misaligned and loose locking nuts causing Error 101, aligned and tightened. Completed system characterization and system is operating properly. | No Injury | 20212406-BV | During Preparation for Use | Yes | 78
167457 | The laser system does not provide reliable performance | Unable to replicate complaint | N/A | Power | Other | Non-Reportable | Closed | Yes | N/A | UNITED KINGDOM (GBR) | 100002 | Whilst we were using the laser machine for a case in lab 1 today I noticed that it seemed to take longer than usual for the machine to go back on standby. Usually when I press the button it goes back onto standby almost straight away so just wanted to check that there isn't a problem with the machine. Many Thanks Abi1. The delay was after treating the patient, so going on standby on removing the laser from the patient. It was a 0.9 catheter and we lasered for 4 minutes and 5 seconds2. The delay was probably only around 5-10 seconds but noticeable as usually the machine goes on to standby immediately3. No delay noticed today - I was unable to calibrate a catheter in the machine but I did use the calibration catheter to calibrate the machine and then switched to ready and standby a few times with no issue noted4. No other issues experienced yesterday and no errors or faults appearedMany ThanksAbigail Threadgold ¦ Cardiac Radiographer | FSE was not able to duplicate the reported problem, system operated properly on arrival. | No Injury | 20212419-BV | During Use in Patient | Yes | 70
167896 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | ITALY (ITA) | 100008 | Error discovered during PLS log revise at start of visit on Wednesday October 14, 2021 with call number 20212406-BV.Unreported Error Code 500, “self test failed” from 7 October 2021(13.13) and 20 sept 2021 (05.28)In both cases, the system showed the error on bootup first thing in the morning, they restarted the system and continued without issue. | Error 500 was found in the log while onsite for complaint #166710. Error due to shutter sticking and was not duplicated during this visit. | No Injury | N/A | During Use in Patient | Yes | 82 | 500
168147 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | GERMANY (DEU) | 100009 | During installation the system displayed an Error 500 on power up, after 2 restarts the system powered up and operated properly. No patients were involved. | Error 500 on power up, cleared with restart. System is operating properly. | No Patient Present | 20212425-BV | During Installation | No | 78 | 500
168275 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | NETHERLANDS (NLD) | 100013 | During incoming inspection in the BV, Alasdar reported the following on power up:2 Error Code 500, cleared on third startupCould not calibrate any catheters( cath cal error 0201 catheter input saturated) | Error code 500 during power up, cleared with restart. System operating properly. | No Patient Present | 20212455_BV | Not during a case | Yes | 77 | 500
168276 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | NETHERLANDS (NLD) | 100012 | During incoming inspection at BV, Alasdar reported the following:1 error code 500 on start-up, cleared on restart. | Error 500 on power up during incoming inspection, cleared with reboot. System is operating properly. | No Patient Present | Non-Dispatch | Not during a case | Yes | 78 | 500
168402 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Bad Clinical Catheter | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100000 | error code 201• Reference catheter works properly.• All three used sterile 2.5 catheters connected to the PLS immediately gave out the Fault 0201. PLS repair & catheter replacement delivery | Catheter is not supported by the PLS system, 2.5mm TE with 0.18 wire. System operating properly. | No Injury | 20212444 | During Preparation for Use | Yes | 73 | 201
170718 |  |  | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No |  | GERMANY (DEU) | 100001 | 106 sensor mismatch error appeared during installation, please refer to the installation call 20212506-BV.Issue was resolved on-site. Please also see correspondence attached as a reference. |  | No Patient Present | 20212506-BV | During Installation |  | 20 | 106
170900 |  |  | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | N/A | UNITED KINGDOM (GBR) | 100002 | FSE responded the e-call 167457 with call number 20212419-BV on 09 November. On December 01, I noticed on one of the emails that the FSE encountered error 106 during the service. |  | No Patient Present | 20212508-BV | Routine Service/Not during a case |  | 15 | 106
172906 | The laser system is not firing/lasing | Fault 106 | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100046 | During the prep for the case, the reference catheter and 0.9mm clinical catheter calibrated properly. The fluence/rate was set to 40/40. When attempting to lase the laser displayed 'System Error'. No code was associated with the error. They opened a new clinical catheter with the same result. Case was not completed with the PLS Nexcimer. They used their CVX to complete the case at the same facility with no delay in treatment. | FSE verified the error 106, adjusted attenuator, completed system characterization and system is operating properly. | No Injury | 20220020 | During Use in Patient | Yes | 82
173042 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100046 | Customer called to report an Error 310 during the procedure, lead management. Case was completed using the CVX on site with no issues. | FSE was not able to duplicate the reported problem, completed calibrations and system is operating properly. | No Injury | 20220048 | During Use in Patient | Yes | 85 | 310
173214 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | SPAIN (ESP) | 100004 | The laser was failing during a procedure with the FC 108. This happens if immediately after pressing Ready the Footswitch is pressed. A rebooting solves the problem. This issue was fixed over the phone. | Non dispatch phone fix, Error 108 is due to pressing ready and the footswitch immediately, system was rebooted and is operating properly. | No Injury | N/A | During Use in Patient | Yes | 91 | 108
173452 | The laser system has loose components | Keyswitch | Unknown | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100017 | Rep called to report the keyswitch is loose and comes out of the rear panel, system powers up and operates properly. Also reported an unknown error on power up that was cleared with a reboot. Case was completed without issue at the same facility. | FSE found a damaged keyswitch assembly, replaced. Unknown fault was error 500, shutter is operational and system is operating properly. | No Injury | 20220241 | During Preparation for Use | Yes | 85
173720 | The laser system has broken components | Electro / Mechanical | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | MN | UNITED STATES (USA) | 100037 | The handle and wheels were damaged in shipping and need to be replaced. | During installation the FSE noted damage to the rear handle and casters, replaced all. Completed calibrations and system is operating properly. | No Patient Present | 20213533 | During Installation | Yes | 83
173947 | The laser system will not turn on | Interlocks Intermittently Dis-Engage | N/A | Alarms/ Messages | Constant alarm | Non-Reportable | Closed | Yes | CT | UNITED STATES (USA) | 100043 | Adam said when he plugged in the PLS Laser it started buzzing/alarm and he could not turn on the laser. I had him reset the Red Estop button and he said it did not correct this Alarm. Case will be completed with their other laser. | FSE found the rear panel interlock was not fully disengaged, tightened the rear panel. Completed courtesy checks and system is operating properly. | No Injury | 20220277 | During Preparation for Use | Yes | 82
173970 | Clinical catheter fails to calibrate | Catheter Coupler Assy | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100023 | During the prep for the case, the reference catheter would not successfully calibrate. They attempted to calibrate the reference catheter four times with the same result. They opened a 0.9mm clinical catheter and it would also not calibrate. They attempted multiple times with the same result. No error message was observed, only that the calibration was unsuccessful. The case was not completed with the PLS. The same catheter that failed on the PLS calibrated with no issues with their CVX300. The case was completed with their CVX300 laser. | FSE found a defective catheter coupler assembly, replaced. Completed calibrations and system is operating properly. | No Injury | 20220281 | During Preparation for Use | Yes | 82
174673 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Catheter Aiming During Calibration | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100030 | During the prep for the case, during calibration, an error 103 occurred. The FSE instructed them to reset the laser and try again. On the second attempt, the catheter calibrated properly and the case was completed with the laser. | Non dispatch phone fix, FSE contacted customer and determined the calibration technique was poor. Instructed to reboot the system, calibrate again, system calibrated and operated properly. | No Injury | N/A | During Preparation for Use | Yes | 80
174957 | Multiple Fault Codes Displayed | Fault 106 | Fault 302 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100008 | Dear Alasdar and dear all,I’m recalling this old conversation related to the issue with the new PLS installed in Perugia.It seems it is not working properly again.In more detail, the PLS was used 2 days ago during lead extraction procedures and it worked properly.Today, we received an email from the clinical engineering department asking for technical assistance because, during a procedure in the interventional cardiology department, the physicians were not able to treat the patient because 2 errors occurred: code 302 and code 106.Can you please help me to better understand what these codes refer to? And if there is something I can do or if we should plan a new technical intervention to solve the issue?Thank you so much in advance,Charlotte GuarrasiClinical Application Specialist Central ItalyImage Guided Therapy DevicesPhilips, Excelsiorlaan 41, 1930 Zaventem, BelgiumEmail charlotte.guarrasi@philips.com Mob +39 347 983 40 33www.philips.com/IGTdevices | FSE found optical rail misalignment and attenuator position issue, aligned and completed zero attenuation and system characterization. System is operating properly. | No Injury | 20220065-BV | During Use in Patient | Yes | 73 | 302
175116 | Multiple Fault Codes Displayed | Fault 106 | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | GREECE (GRC) | 100014 | Hoi Diana,I hope you are doing well!I got several calls from Ducasco our distributor in Greece about the malfunctioning of the PLS.Dr. Koustas tries several times to use the system and it gives some errors specially with the 0.9 & 1,4 TE So the question is when can an Engineer passerby? The next urgent case is planned for Wednesday 23rd noonThey proposed to check the system the same day (morning) in order to performance the case after.Please give me some signs.Thanks in advanceCesar | FSE found the attenuator blind misadjusted, adjusted and completed calibrations. System is operating properly. | No Patient Present | 20220071-BV | Routine Service/Not during a case | Yes | 69
175242 | Clinical catheter fails to calibrate | Dirty Energy Detector | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | N/A | SPAIN (ESP) | 100004 | During prepatarion for use the laser was failing during the calibration. Due to the customer information with a 2.0 non sterile cath. With another catheter the system calibrated well. This is not very clear what catheter was used as the reference cath is a 2.5 cath. There is a sheduled service call on the 16.2.2022 and the system will be checked. So far this issue was solved over the phoneThe intervention with the laser went well with a sterile cathAbout the kind of treatment are no details available | FSE found debris on the face of the catheter detector, cleaned and completed system calibrations, operating properly. | No Injury | 20220032-BV | During Preparation for Use | Yes | 68
175388 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100016 | Don said they got a 106 error code, reset the laser and was able to finish the case. Patient is doing well. | FSE was not able to duplicate the reported problem, completed system characterization and noted slight sensor drift. System is operating properly. | No Injury | 20220522 | During Use in Patient | Yes | 74 | 106
175604 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100017 | Rep emailed to report an Error 230 at power up, CVX was used to complete the case. | FSE was not able to duplicate the reported problem, system powered up and operated properly on arrival. | No Injury | 20220542 | During Preparation for Use | Yes | 62 | 230
175621 | The laser system does not provide reliable performance | Fault 108 | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100045 | Customer called to report an unknown "red error" issue during the procedure, system was "reset" but unable to determine if the case was completed with the laser system. There was no harm to the patient. FSE will attempt to gather additional information once on site. | FSE determined the unknown error was 108, mode state change error, completed training for the customer. | No Injury | 20220543 | During Use in Patient | Yes | 64
175940 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | HI | UNITED STATES (USA) | 100027 | During the case, an error 108 occurred. Laser was reset and the case was completed with the laser with no further issues. Error 108 is caused by pressing on the footswitch too soon after going into the 'Ready' mode. FSE to contact account for training / phone fix. | Non dispatch phone fix, error due to operator dressing footswitch immediately after going into ready mode. FSE contacted customer and distributor for training. | No Injury | N/A | During Use in Patient | Yes | 66 | 108
176197 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | Error 106. Case was not completed with the laser, it was completed with a CSI device. | FSE found optical rail alignment issue, aligned, completed sensor and syschar calibrations and system is operating properly. | No Injury | N/A | During Use in Patient | Yes | 90 | 106
176305 | The footswitch connection is not functioning | Fault 251 | Fault 251 | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100042 | During the prep for the case, after calibration, the footswitch disconnected from the rear of the laser and moved to the front. The laser did not recognize that the footswitch was connected to the front connector, Exception 251. The laser had to be restarted and the laser recognized the footswitch at the front connector. Case was completed with the laser. | FSE was not able to duplicate the reported problem, footswitch connected in front and rear connections without issue. | No Injury | 20220600 | During Preparation for Use | Yes | 59 | 251
176650 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100071 | Fault Code 500/501 at start up | FSE was not able to duplicate the reported problem but issue was verified in the log. Shutter was cycled several times without issue, completed calibration and system is operating properly. | No Patient Present | 20220625 | Not during a case | Yes | 63 | 500
177173 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100059 | Cedric (FSE) said the laser was giving a Fault 106 and they could not complete the case with the laser.The case was completed with AngioSculpt, Patient is doing well | FSE found a defective attenuator, replaced. completed calibrations and system is operating properly. | No Injury | 20220922 | During Use in Patient | Yes | 62 | 106
177661 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100041 | Jim called to report the system calibrates the reference catheter but failed to calibrate a 0.9mm catheter, no error codes displayed. Another catheter was opened, calibrated, and the case was completed without issue. The failed catheter will be available on site for testing, troubleshooting. | FSE found optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20220838 | During Preparation for Use | Yes | 48
177801 | The laser system does not provide reliable performance | Rare Occurrence | N/A | Software/ Firmware (excluding privacy or security) | Date/Time issue | Non-Reportable | Closed | Yes |  | UNITED KINGDOM (GBR) | 100002 | A courtesy call was carried out in order to determine if the date and time stamp had reverted back to factory default as this was discovered on the last service to be the case. They had indeed reverted back and a plan is being implemented on how best to rectify the issue. Where as this issue does not affect the systems functionality it does hinder any future investifations that may need to be carried out as the date in the logs resets every time the system is rebooted. | FSE noted during a courtesy check that the internal clock resets to 1970 on power up, Replaced the HMI display panel. Completed calibrations and system is operating properly. | No Patient Present | 20220161-BV | Routine Service/Not during a case | Yes | 44
178009 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100031 | Error 106 after calibration. PLS was restarted with the same result. Case was not completed with the laser. Unknown how the case was completed. Sales Rep Marisa Boling will contact facility concerning how the case was completed. | FSE was able to duplicate the E106. PVI was completed and system is operating properly. | No Injury | 20220881 | During Use in Patient | Yes | 49 | 106
178072 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100061 | E106 and E108 while the Sales Rep and doctor were modeling and training with the laser. The error was persistent with two ELCA 0.9mm catheters. No patient involvement. | FSE was unable to duplicate the error 106. PVI completed and system is operating properly. | No Patient Present | 20220886 | Not during a case | Yes | 46 | 106
178576 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100059 | Contact: Sales Rep Trevis Bachmann @ 480-282-8080Issue: Laser gave a FC 106 – they had to go thru 2 catheters: 1.4 gave 106 error code. The 1.7 was used next and that did no calibrate,then a 0.9 catheter was used and it calibrated and worked. It was a good outcome.This happened today, during a case. Case completed with the laserTrevis did not observe the case. Dr. Preshad observed the case. | FSE completed sensor cal, system characterization, and alignment, system is operating properly. | No Injury | 20220922 | During Use in Patient | Yes | 40 | 106
178696 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Unable to replicate complaint | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100024 | Laser gave 106 error | FSE was not able to duplicate the reported problem, completed alignment and calibrations system is operating properly. | No Injury | 20220925 | During Preparation for Use | Yes | 39 | 106
178751 |  |  | Fault 108 |  |  |  | Closed | No | VA | UNITED STATES (USA) | 100042 | Jamie said they got a fault code 108 during the lead removal. I asked her to reboot the power to the laser and recalibrate the catheter. Waiting to hear from Jamie if this worked or not and if they completed the case with the laser? |  | No Injury | N/A | During Use in Patient |  | 0 | 108
178752 |  |  | Fault 108 |  |  |  | Closed | No | VA | UNITED STATES (USA) | 100042 | Jamie said they got a fault code 108 during the lead removal. I asked her to reboot the power to the laser and recalibrate the catheter. Waiting to hear from Jamie if this worked or not and if they completed the case with the laser? |  | No Injury | N/A | During Use in Patient |  | 0 | 108
178753 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100042 | Jamie said they got a fault code 108 during the lead removal. I asked her to reboot the power to the laser and recalibrate the catheter. Waiting to hear from Jamie if this worked or not and if they completed the case with the laser? Jamie said the case was completed with their other laser the CVX300. Patient is doing well. Doc is not happy with the PLS Laser. | Non dispatch phone fix. FSE confirmed the customer was depressing the footswitch too quickly after depressing ready. System is operating properly. | No Injury | 20220929 | During Use in Patient | Yes | 36 | 108
179123 | The laser system is not firing/lasing | Fault 106 | N/A | Calibration | Failed to maintain calibration | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100033 | Tara called in an issue with the PLS SN 100033. The turbo power catheter they were using would not rotate when plugged into the PLS, but had no issues when plugged into the CVX system. Account used their other laser to complete the case. Patient is doing well. | FSE could not duplicate the reported problem, review of the log indicated a 106 error, completed calibrations and system is operating properly. | No Injury | 20221141 | During Preparation for Use | Yes | 50
179256 | Clinical catheter fails to calibrate | Fault 106 | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100084 | Customer called FSE to report an unknown Error during calibration, case was completed with the CVX system onsite. | FSE duplicated an Error 106, completed sensor and system calibrations however problem persisted. System will be deinstalled and replaced with SN 100101. | No Injury | 20221161 | During Preparation for Use | Yes | 52
179363 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100086 | Customer called to report an E106 with a 1.7mm catheter when running at a setting of 60/80. Customer rebooted the system, recalibrated, and completed the case at a lower setting. | FSE verified E106, completed sensor and system calibration. System is operating properly. | No Injury | 20221171 | During Use in Patient | Yes | 56 | 106
179512 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 104 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100081 | Customer called to report an Error 104, the power was cycled twice and produced the same error after calibration. Case was completed by other means at the same facility. | FSE verified an E106 misreported as 104, completed sensor and system calibrations and system is operating properly. | No Injury | 20221180 | During Use in Patient | Yes | 55 | 104
179518 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100086 | E106 after calibration of 1.7mm Turbo. Restarted laser and completed the case with the laser at 55 fluence and 75 rate.Additional complaint for 106 noted, FSE on site for 106 reported on 4/19, CID#179363. | See service report 20221171 and CID #179363 for summary. Same failure the following day, completed within the same service report. | No Injury | 20221171 | During Use in Patient | Yes | 43 | 106
179951 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100081 | Error 106 with a 0.9mm catheter at 45/45. Laser ran for 30 seconds and then gave the error. Case was not completed with the laser. Case was completed with a scoring balloon and drug coated balloon. | FSE verified E106, completed sensor and system calibration and system is operating properly. | No Injury | 20221204 | During Use in Patient | Yes | 71 | 106
179968 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100033 | E108. Advised to allow a delay between pressing the 'Ready" and stepping on the foot switch. Continuing the case with the laser. | Tech support determined the customer is depressing the footswitch within 2 seconds of depressing ready, system is operating properly. | No Injury | N/A | During Preparation for Use | Yes | 43 | 108
180042 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100035 | E106 with a 0.9mm catheter. Case was not completed with the laser. Sales Rep to confirm on 02Apr2022 how the case was completed.02May2022: FSE reported that the customer completed the case with the laser at the default Fl & Rate. | FSE verified reported problem, noted alignment was slightly off and aligned. Completed calibrations and system is operating properly. | No Injury | N/A | During Use in Patient | Yes | 45 | 106
180198 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100086 | Amy said they got a fault code 106, They lowered the rate and fluency 55/65 and was able to complete the case with the laser. Patient is doing well. | FSE found VES defective, replaced. Completed calibrations and system is operating properly. | No Injury | 20221222 | During Use in Patient | Yes | 44 | 106
180269 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100051 | Luke Meissner (rep) called in error 106 for SN 100051. Unable to complete case with PLS, used CSI atherectomy device to complete the case. Patient is doing well. | FSE verified E106, adjusted alignment and sensor calibration, completed system calibrations and operating properly. | No Injury | 20221229 | During Use in Patient | Yes | 43 | 106
180563 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IN | UNITED STATES (USA) | 100099 | clinical rep said that they had an E106 error on their first case today. 2.0 Catheter run at 50/40. | FSE verified E106 and completed attenuator adjustment. Completed calibrations and system is operating properly. | No Injury | 20221256 | During Use in Patient | Yes | 51 | 106
180753 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100085 | Laser gave E106. | FSE verified E106, completed alignment and calibrations, system is operating properly. | No Injury | 20221269 | During Preparation for Use | Yes | 46 | 106
180906 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 105 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Arizona | UNITED STATES (USA) | 100083 | Rep called to report an Error 105 after calibration with a 2.0 at 60/80. Rebooted the system and was able to complete the case. | FSE found optical rail efficiency above specification, completed calibrations and system is operating properly. | No Injury | N/A | During Use in Patient | Yes | 53 | 105
181044 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100060 | Rep called to report an E106 with a .9 catheter at 60/80. Case was completed with CSI device. | FSE verified E106, completed sensor and system calibration, system is operating properly. | No Injury | 20221464 | During Use in Patient | Yes | 50 | 106
181137 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100019 | Together with a proctor, coming from another hospital and a clinical application specialist I was supporting a lead extraction procedure.The patient was ready to be lasered with 2 PM Leads from the right side. The 2 LLD`s were in place. I tried since 10 minutes to calibrate the PLS, but after switching the key I received the warning: Systemfehler 500, kontaktieren sie den Kundendienst. I tried again and again, switched the machine on and off, and received the same warning every time. I tried another connector, again the same warning. I tried to reach our field service engineers. Finally I reached Alasdar. He told me to try it several times more until it works. After switching the machine about 16 times on and off it finally worked, I could calibrate and continue the procedure as planned. BUT the physicians were scared and also the whole team preparing to use TightRail because the completely new PLS did not work as it should. The PLS was installed in Feb.2022 and now it was the first use. | Tech support determined the shutter was stuck on the bumper and cleared, system is operating properly. | No Injury | N/A | During Preparation for Use | Yes | 50 | 500
181291 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100085 | Rep called to report an E106 with a .9mm device at 80/80. Case was completed by other means at the same facility. | FSE verified E106, found defective VES and PES sensors, replaced. Completed calibrations and system is operating properly. | No Injury | N/A | During Use in Patient | Yes | 54 | 106
181809 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100015 | Rep called to report an E108, rebooted system and completed case without issue. Caused by depressing the footswitch too quickly after selecting ready. Instructed to wait a couple of seconds after pressing ready to depress footswitch. | Tech support instructed customer to wait a few seconds after pressing ready before depressing footswitch. System is operating properly. | No Injury | N/A | During Preparation for Use | Yes | 44 | 108
182439 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100031 | Marissa called in Fault 106. Able to complete case using cutting balloon and finished with balloon angioplasty. Patient is doing well | FSE was not able to duplicate the E106, completed calibrations and system is operating properly. | No Injury | 20221572 | During Preparation for Use | Yes | 49 | 106
182505 | Clinical catheter fails to calibrate | Fault 500 | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100026 | David said while trying to calibrate the PLS laser they got an unsuccessful try message several times. They ended up bringing in the other Laser (CVX300) and prepared it to be used but the lead removal did not require the laser as it came out easily. Patient is doing well. | FSE was not able to duplicate the reported problem, review of log indicated an E500. Completed system calibrations and system is operating properly. | No Injury | 20221585 | During Preparation for Use | Yes | 48
182506 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Co | UNITED STATES (USA) | 100075 | Kelly said they got the fault 106 during the case. They had made several passes with laser before it faulted out. Case was completed with Balloon and Stent. | FSE verified E106, completed sensor and system calibrations. System is operating properly. | No Injury | 20221575 | During Use in Patient | Yes | 48 | 106
182521 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100033 | Wade said they got a Fault 106 during the case. case was completed with Balloon. Patient is doing well. | FSE was not able to duplicate the reported problem, E106 was recorded in the log file. Completed calibrations and system is operating properly. | No Injury | 20221577 | During Use in Patient | Yes | 48 | 106
182657 |  |  | Fault 500 |  |  |  | Closed | No |  | GERMANY (DEU) | 100018 | The customer is complaining that this laser is failing frequently with the faultcodes 500, 511 and 211. After rebooting he was able to comlete all cases. We recommended the customer to reboot by troubleshooting. this interrupts the operations massively, but currently no patient unhurted. For this service we need urgently the new shutter assembly.CS notes: after the reboot of the system the Laser was up and running again and they completely finish the intervention. Patient is safe and no injury. On June 14, 2022 our FSE will visit the system to repalce the shutter. |  | No Injury | 20220274-BV | During Use in Patient |  | 0 | 500
182658 | Multiple Fault Codes Displayed | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100018 | The customer is complaining that this laser is failing frequently with the fault codes 500, 501 and 211. After rebooting he was able to compete all cases. We recommended the customer to reboot by troubleshooting. this interrupts the operations massively, but currently no patient unhurted. For this service we need urgently the new shutter assembly.CS notes: after the reboot of the system the Laser was up and running again and they completely finish the intervention. Patient is safe and no injury. On June 14, 2022 our FSE will visit the system to replace the shutter.TS notes: Review of the log file indicates exceptions occurred on both the 2nd of June and the 9th of June, all about the same time frame. E500 - boot up error, E501 - boot up timeout, E211 - Shutter closed. All related to the shutter and repaired per CR156370. Requested a complaint be generated for the June 2nd event through tech support, BV. | FSE verified E500, replaced the shutter assembly under CR156370. Completed calibrations and system is operating properly. | No Injury | 20220274-BV | During Use in Patient | Yes | 51 | 500
182892 | The laser system has broken components | Emergency Switch | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Emergency Stop button is broken and fell off. Case completed with laser | FSE verified damaged emergency switch, replaced. Completed calibrations and system is operating properly. | No Injury | 20221776 | During Preparation for Use | Yes | 53
182902 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100046 | Error found in log found during routine maintenance. This error did not happen during service.Error 252 FOOTSWITCH_INVALID quant. 1 | During routine service the FSE noted an E252 in the log, system operated properly throughout the service. | No Patient Present | 20221430 | Routine Service/Not during a case | Yes | 50 | 252
182903 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100046 | Error found in log found during routine maintenance. This error did not happen during service.Error 108 MAX_INVALID_PULSES quant. 1 | During routine service the FSE noted an E108 in the log file, system operated properly throughout the service. | No Patient Present | 20221430 | Routine Service/Not during a case | Yes | 50 | 108
182905 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Unable to replicate complaint | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100046 | Error found in log found during routine maintenance. This error did not happen during service.Error 103 CATHETER_INPUT_SATURATED quant. 1 | During routine service FSE noted E103 in the log, system operated properly throughout the service. | No Patient Present | 20221430 | Routine Service/Not during a case | Yes | 50
182971 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 105 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100061 | Code 105 was generated. This was found on log file during service of complaint #178072 | During service the FSE noted an E105 in the log, was not duplicated during service. Completed calibrations and system is operation properly. | No Patient Present | N/A | Routine Service/Not during a case | Yes | 50 | 105
182972 |  |  | Fault 106 |  |  |  | Closed | No | MA | UNITED STATES (USA) | 100061 | Code 106 was generated in log file. This was found during service on complaint 178072 |  | No Injury | N/A | Unknown |  | 1 | 106
182974 |  |  | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | MA | UNITED STATES (USA) | 100061 | Code 108 was generated on log file. This was found during service on complaint 178072 |  | No Injury | N/A | Unknown |  | 1 | 108
182976 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 211 | Fault 211 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 1000017 | Code 211 was generated on file log. This was found during service call on complaint 175604 | E211 was found in the log during service, no issues found, completed calibrations and system is operating properly. | No Injury | N/A | During Preparation for Use | Yes | 50 | 211
183115 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 240 | Fault 240 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | SPAIN (ESP) | 100004 | upon reviewing the PLS log, the FSE found out that there was na error of 240 erroro happened on 01 December 2021. The log was pulled out on 15th of Feb 2022. | During service FSE noted an E240 in the log, no issues found, completed calibrations and system is operating properly. | No Patient Present | N/A | Not during a case | Yes | 46 | 240
183241 | Multiple Fault Codes Displayed | Shutter Assembly | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100018 | CS Notes: Reporting this complaint after pulling out the logs of the PLS and found out there was an error E500, 501, 211 on June 02, 2022.This is an additional complaint to #182658 | Please refer to complaint ID 182658. This complaint was noted after reviewing the log file. | No Patient Present | N/A | Not during a case | Yes | 44 | 500
183299 |  |  | Fault 106 |  |  |  | Closed | No | TN | UNITED STATES (USA) | 100023 | Error Code 106 found during the review of the System Log. |  | No Patient Present | 20220281 | Routine Service/Not during a case |  | 5 | 106
183302 |  |  | Fault 106 |  |  |  | Closed | No | Louisiana | UNITED STATES (USA) | 100015 | Error Code 106, 108, 201, 230, 252, 310 found during the review of the System Log. |  | No Patient Present | 20220703 | Routine Service/Not during a case |  | 5 | 106
183303 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100015 | Error Code 106, 108, 201, 230, 252, 310 found during the review of the System Log. | During routine service the FSE noted several errors in the log, no issues were duplicated, completed calibrations and system is operating properly. | No Patient Present | 20220703 | Routine Service/Not during a case | Yes | 50 | 106
183308 | Multiple Fault Codes Displayed | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100026 | Error codes found during review of the System Log.Error:500 self test failedError:500 Self test failedError 501 Power up timeoutError 501:Power up timeout | During service the FSE noted an E500 in the log, no issues duplicated, completed calibrations and system is operating properly. | No Patient Present | 20221585 | Routine Service/Not during a case | Yes | 50 | 500
183425 |  |  | Fault 102 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | AZ | UNITED STATES (USA) | 100028 | Error Codes NL102, NL103, NL104, NL105 found after review of System Log |  | No Patient Present | 20221314 | Routine Service/Not during a case |  | 5 | 102
183430 | Multiple Fault Codes Displayed | Fault 211 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100031 | Error code 211 was found during review of System Log. | During service the FSE noted several error in the log, service was for an E106 that was not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20221572 | Routine Service/Not during a case | Yes | 53 | 106
183562 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Unable to replicate complaint | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | During the morning check of the laser, A FC200 occurred during the self check. No patient involvement. | FSE was not able to duplicate the reported problem, replaced the coupler assembly as a precaution. Completed calibrations and system is operating properly. | No Patient Present | 20221820 | Not during a case | Yes | 53
183576 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Code 106, 108, 221, 252, 301, 310 found during the review of the System Log | During servicing the FSE noted several error codes- 106, 108, 221, 252, 301, 310. No errors were duplicated during servicing, completed calibrations and system is operating properly. | No Patient Present | 20221776 | Not during a case | Yes | 60 | 106
183625 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | SPAIN (ESP) | 100004 | Good morning,I attach the ruling that the laser have yesterday in the H. Valle Hebron.Yesterday a patient underwent surgery, he was already asleep and the procedure was being carried out. The laser was calibrated and then calibrated again with the sterile sheath and all went well. The first cable began to be removed and there was no problem, but with the second cable the laser began to fail 106. This led to stopping the surgery, turning the laser off and on again. It calibrated well, but when operating, error 106 appeared again.In the end, it was operated with mechanical devices. Full extraction.On Wednesday there will be another surgery and next week two more, can Olaf go check it out?Thanks for your help.Kind RegardsElisa. | FSE verified E106 in the log but was not able to duplicate, completed alignments and calibrations and system is operating properly. | No Injury | 20220343-BV | During Use in Patient | Yes | 59 | 106
183785 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100028 | Error code 106 found during review of system log. | During routine service the FSE noted an E106 in the error log, was not able to duplicate the issue, completed calibrations and system is operating properly. | No Patient Present | 20221314 | Routine Service/Not during a case | Yes | 78 | 106
183787 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100059 | Error code Exception 221 ‘Attenuator position Fault’ found during review of system log. | During service the FSE noted an E221 in the log, was not able to duplicate, replaced attenuator and completed calibrations. System is operating properly. | No Patient Present | 20220820 | Routine Service/Not during a case | Yes | 82 | 221
183789 |  |  | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | CA | UNITED STATES (USA) | 100030 | Error code 103 found during review of system log. |  | No Patient Present | 20221365 | Routine Service/Not during a case |  | 1 | 103
184058 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | 30 Day Reportable | Closed | Yes |  | SPAIN (ESP) | 100004 | Dear all,Right now a interventional cardiology doctor is working in H. Valle Hebron (with a patient on the table) and the PLS is giving a “calibration fault” as show the attached video.They have already oppened 2 elca catheters and anyone works well.There is anyone that can help us now?Is there something we can do now for arrange that?Many thanks to all of you.Additional information was received from the distributor via email on 11 July 2022:1) Since the laser system could not start up, what equipment/device(s) were used to treat the patient? The parient has been postponed.2) Was the patient treated the same day at the same hospital? Or was there a delay in treatment?A delay in the treatment. 3) At the hospital, were there alternate means of treatment, since the laser system did not start up?They have alternatives for the treatment but due the special features of the patient they want to treat him with the laser.4) Did the patient require transfer to another hospital to receive treatment?Right now the patient has not been transfer to another hospital. The doctors are waiting until tomorrow to see if the laser is repared. If not, maybe they will consider to transfer him. | FSE was not able to duplicate the reported problem, system operated properly on arrival. FSE did note the attenuator was mounted on the incorrect location of the motor shaft, adjusted, completed attenuator zero routine and calibrations. System is operating properly. | No Injury | 20220358-BV | During Use in Patient | Yes | 45
184128 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | HI | UNITED STATES (USA) | 100027 | Error Code 500 found during the review of the System Log | During routine service the FSE noted an E500, not duplicated, system operated properly. | No Patient Present | 20221328 | Routine Service/Not during a case | Yes | 76 | 500
184129 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 501 | Fault 501 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | HI | UNITED STATES (USA) | 100027 | Error Code 501 found during the review of the System Log | During routine service the FSE noted an E501, not duplicated, system operated properly. | No Patient Present | 20221328 | Routine Service/Not during a case | Yes | 76 | 501
184131 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | HI | UNITED STATES (USA) | 100027 | Error Code 108 found during the review of the System Log | During routine service the FSE noted an E108, not duplicated, system operated properly. | No Patient Present | 20221328 | Routine Service/Not during a case | Yes | 76 | 108
184133 | Multiple Fault Codes Displayed | Fault 231 | Fault 303 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | HI | UNITED STATES (USA) | 100027 | Error Codes 231, 240 & 303 found during the review of the System Log | During routine service the FSE noted E231, E240, E303, not duplicated, system operated properly. | No Patient Present | 20221328 | Routine Service/Not during a case | Yes | 76 | 303
184193 | The laser system will not turn on | Damaged Control Panel | N/A | Power | Will not turn on | Non-Reportable | Closed | Yes |  | ISRAEL (ISR) | 100072 | New PLS did not work after installation on 05/07/2022 Igor Jerbic is the engineer on site performing this.it never have been used before. | During installation the FSE found the system HMI would not boot up, replaced. Completed calibrations and the system is operating properly. | No Patient Present | 20220353-BV | Routine Service/Not during a case | Yes | 75
184246 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 201 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100034 | Error Code 201 found during the review of the System Log. | During routine service the FSE noted an E201, not duplicated, system operated properly. | No Patient Present | 20221358 | Routine Service/Not during a case | Yes | 78 | 201
184249 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100034 | Error Code 108 found during the review of the System Log. | During routine service the FSE noted an E108, not duplicated, system operated properly. | No Patient Present | 20221358 | Routine Service/Not during a case | Yes | 75 | 108
184250 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100034 | Error Code 108 found during the review of the System Log. | During routine service the FSE noted E108 in the log, not duplicated. System is operating properly. | No Patient Present | 20221358 | Routine Service/Not during a case | Yes | 75 | 108
184255 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log. | During routine service the FSE noted E108 in the log, not duplicated. System is operating properly. | No Patient Present | 20221429 | Routine Service/Not during a case | Yes | 75 | 108
184258 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log. | During routine service the FSE noted E108 in the log, not duplicated. System is operating properly. | No Patient Present | 20221429 | Routine Service/Not during a case | Yes | 75 | 108
184260 | Multiple Fault Codes Displayed | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Codes 108 & 301 found during the review of the System Log. | During routine service the FSE noted E108 in the log, not duplicated. System is operating properly. | No Patient Present | 20221429 | Routine Service/Not during a case | Yes | 75 | 108
184262 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log. | During routine service the FSE noted E108 in the log, not duplicated. System is operating properly. | No Patient Present | 20221429 | Routine Service/Not during a case | Yes | 75 | 108
184263 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log. | During routine service the FSE noted E108 in the log, not duplicated. System is operating properly. | No Patient Present | 20221429 | Routine Service/Not during a case | Yes | 75 | 108
184265 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Rare Occurrence | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 103 found during the review of the System Log. | During routine service the FSE noted E103 in the log, not duplicated. System is operating properly. | No Patient Present | 20221429 | Routine Service/Not during a case | Yes | 75 | 103
184266 | Multiple Fault Codes Displayed | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Codes 252 & 303 found during the review of the System Log. | During routine service the FSE noted E252 in the log, not duplicated. System is operating properly. | No Patient Present | 20221429 | Routine Service/Not during a case | Yes | 75 | 252
184312 |  |  | Fault 103 | Alarms/ Messages |  |  | Closed | No | TN | UNITED STATES (USA) | 100046 | Error Code 103 found during the review of the System Log. |  | No Patient Present | 20221430 | Routine Service/Not during a case |  | 0 | 103
184314 |  |  | Fault 108 |  |  |  | Closed | No | TN | UNITED STATES (USA) | 100046 | Error Code 108 found during the review of the System Log. |  | No Patient Present | 20221430 | Routine Service/Not during a case |  | 0 | 108
184315 |  |  | Fault 252 |  |  |  | Closed | No | TN | UNITED STATES (USA) | 100046 | Error Code 252 found during the review of the System Log. |  | No Patient Present | 20221430 | Routine Service/Not during a case |  | 0 | 252
184450 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100023 | Error Code 108 found during the review of the System Log. | During routine service the FSE noted an E108 in the log, not duplicated. System is operating properly. | No Patient Present | 20221768 | Routine Service/Not during a case | Yes | 73 | 108
184451 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100023 | Error Code 252 found during the review of the System Log. | During routine service the FSE noted an E252 in the log, not duplicated. System is operating properly. | No Patient Present | 20221768 | Routine Service/Not during a case | Yes | 73 | 252
184554 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100031 | E106 reported. The error was reset from cycling power and a different catheter was used without any issue. | FSE verified E106 with a .9 catheter and was unable to correct the issue with alignment and calibrations. The system will be deinstalled and replaced with the next available system. | No Injury | 20221879 | During Use in Patient | Yes | 72 | 106
184814 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Loose Hardware | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100073 | Andre said they were getting a 500 and 501 error code on their PLS laser , I had him reset power to the laser several times but was unable to get past the errors, Case completed with Balloon and Angioplasty.Patient is doing well. Please call Andre with an eta for service. Andre Henderson @ 203-243-9847 | FSE verified E500, reseated connection for the attenuator on the SCM PCB. Also replaced the shutter assembly per CR156370. Completed calibrations and system is operating properly. | No Injury | 20221903 | During Preparation for Use | Yes | 75 | 500
184841 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100080 | Jeff said they got a 106 error code , reduced the laser output energy/rate and was able to complete the case with the laser. Patient is doing well. | FSE could not verify E106, completed calibrations and system is operating properly. | No Injury | 20221905 | During Use in Patient | Yes | 81 | 106
184849 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100044 | Error codes 106, 221, and 302 found after review of system log | During routine service the FSE noted multiple exceptions in the log not duplicated during service. Completed calibrations and the system is operating properly. | No Patient Present | 20221735 | Routine Service/Not during a case | Yes | 75 | 106
184854 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100047 | Error codes 106 and 108 found after review of system log | During routine service FSE noted Fault 106 and 108 in the log, not duplicated during service. Completed calibrations and system is operating properly. | No Patient Present | CRN 20221736 | Routine Service/Not during a case | Yes | 82 | 106
184991 | Multiple Fault Codes Displayed | Fault 201 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100080 | Error codes 201, 240 & 230 found during review of system log. | During services the FSE noted Error 201, 240, 230 in the log, not duplicated during services. Completed calibrations and system is operating properly. | No Patient Present | 20221905 | Routine Service/Not during a case | Yes | 80 | 201
185037 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100063 | Rep called to report an E106, case was completed by other means at the same time and facility. | FSE verified reported problem, found debris on the 45 mirror and completed attenuator zero. Completed calibrations and system is operating properly. | No Injury | 20222079 | During Use in Patient | Yes | 78 | 106
185309 | Multiple Fault Codes Displayed | Fault 210 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | INDIA (IND) | 100076 | Error 210 recorded in logs on 29th March2022.Error 500 & 501 found reoccurring in this system and was solved by multiple restarts. | FSE was not able to duplicate the reported problems found in the log during service. System is operating properly. | No Injury | N/A | During Preparation for Use | Yes | 77 | 500
185310 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | INDIA (IND) | 100076 | Error 106 while performing procedure with 0.9mm Catheter. Customer was able to calibrate catheter but unable to complete LASING due to error 106 | FSE verified an E106 during service, resolved with zero attenuator and calibration routines. System is operating properly. | No Injury | 20222138 | During Use in Patient | Yes | 77 | 106
185636 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CO | UNITED STATES (USA) | 100070 | When performing in service for staff laser gave E106. | FSE was unable to verify E106, completed alignments and calibrations, system is operating properly. | No Patient Present | 20222108 | Not during a case | Yes | 71 | 106
185748 | The laser system has malfunctioning components | Damaged Control Panel | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100055 | During PM service found that HMI is not saving current time. This problem does not allow log files to be saved to USB drive. | During routine service the FSE noted the system indicated the incorrect time, found defective HMI and replaced. Completed calibrations and system is operating properly. | No Patient Present | 20221644 | Routine Service/Not during a case | Yes | 82
185776 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Optical Rail Mis-Aligned | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100055 | I received a call from Coronary Territory Manager, David Rough. David states the customer at Newark Beth Israel Hospital had problems with 0.9 catheters. 1st catheter failed with error 103. Second catheter calibrated and ran, but eventually also received error 103. The case was completed with the laser and the patient is ok. I’m scheduled to go back to this facility tomorrow to replace the control panel due to internal clock not saving current time. I’ll check the laser and both catheters that failed today. If you have questions, please reply to this email; David or I can answer any questions. | FSE found optical rail was misaligned, aligned and completed calibrations. System is operating properly. | No Injury | 20221644 | During Use in Patient | Yes | 82 | 103
185901 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 240 | Fault 240 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Error Code 240 found during the review of the System Log. | During the routine service the FSE noted an E240 in the service log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222061 | Routine Service/Not during a case | Yes | 81 | 240
185902 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Error Code 108 found during the review of the System Log. | During a routine service the FSE noted an E108, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222061 | Routine Service/Not during a case | Yes | 81 | 108
185942 | The laser system exhausts excessive heat | Unable to replicate complaint | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100079 | During the self-test, the laser had a "Laser too hot" message. Case not completed with the laser. Case completed with Tightrail. | FSE was not able to duplicate the reported problem, noted E600 in the log. Completed calibrations and system is operating properly. | No Injury | 20222116 | During Preparation for Use | Yes | 100
186469 |  |  | Fault 108 |  |  |  | Closed | No | KS | UNITED STATES (USA) | 100067 | Error 108 was found during reveiw of system log |  | No Patient Present | 20222027 | Routine Service/Not during a case |  | 1 | 108
186470 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | Error 108 was found during reveiw of system log | During routine service the FSE noted an E108 in the log not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222027 | Routine Service/Not during a case | Yes | 88 | 108
186471 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Tamil Nadu | INDIA (IND) | 100095 | Error 106 while performing procedure. | FSE found moisture on the PES sensor and prism, cleaned. Completed calibrations and system is operating properly. | No Injury | 20222148 | During Use in Patient | Yes | 88 | 106
186619 | The footswitch cable or cover is damaged | Electro / Mechanical | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100100 | Rep called to report the front footswitch connection is misaligned with the front enclosure. The rear footswitch connection was used to complete the case without issue. | FSE verified the footswitch receptacle misaligned with the outer cover, adjusted. Completed calibrations and system is operating properly. | No Injury | 20222159 | During Preparation for Use | Yes | 86
186736 | Multiple Fault Codes Displayed | Fault 108 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IL | UNITED STATES (USA) | 100054 | Errors 200,201,240,108,105 found after review of system log. | During routine service the FSE found multiple errors in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | CRN 20222134 | Routine Service/Not during a case | Yes | 85 | 201
186740 |  |  | Fault 500 |  |  |  | Closed | No | MN | UNITED STATES (USA) | 100039 | Error codes 500 and 501 found during review of system log. |  | No Patient Present | N/A | Routine Service/Not during a case |  | 1 | 500
186741 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MN | UNITED STATES (USA) | 100039 | Error codes 500 and 501 found during review of system log. | During routine service the FSE noted E500/E501 in the logs, completed shutter upgrade. Completed calibrations and and system is operating properly. | No Patient Present | N/A | Routine Service/Not during a case | Yes | 85 | 500
186873 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | SPAIN (ESP) | 100021 | The fault 500 happened during the prepartion for use. The system was rebooted a few times which did solve the problem. A PM Service is already planned in 2 weeks and my part is available (Shutter) to solve this issue permanently. | FSE verified E500 and replaced the shutter per update. Completed calibrations and system is operating properly. | No Injury | 20220421-BV | During Preparation for Use | Yes | 81 | 500
186889 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CO | UNITED STATES (USA) | 100070 | Allison the TM said they got an Fault 106 while using a 14 french catheter, Recalibrated the catheter and was able to use the laser and complete the case. Patient is doing well. | FSE found optics misaligned and adjustments loose, tightened and aligned. Completed calibrations and system is operating properly. | No Injury | 20222328 | During Use in Patient | Yes | 81 | 106
187087 | Multiple Fault Codes Displayed | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100060 | Error 108 and 252 found after review of system log. | During routine service the FSE noted 108 and 252 in the system log, not duplicated, completed calibrations and system is operating properly. | No Patient Present | 20222036 | Routine Service/Not during a case | Yes | 83 | 108
187137 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100033 | Wade Wegescheide @ 314-783-1160 said the customer got a 106 Error during the case. Case was comepleted with Rotorblade product. patient is doing well. | FSE verified E106 in the log but not duplicated during service, completed calibrations and system is operating properly. | No Injury | 20222342 | During Use in Patient | Yes | 82 | 106
187455 | Clinical catheter fails to calibrate | Electro / Mechanical | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100023 | Multiple attempts to calibrate laser failed. Message of "unsuccessful calibration". Case completed with balloon | FSE found issue with CES, replaced. Completed alignments and calibrations and system is operating properly. | No Injury | 20222372 | During Preparation for Use | Yes | 78
187590 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100059 | E106 when trying to use the laser | FSE verified E106, completed calibrations and system is operating properly. | No Injury | 20222380 | During Preparation for Use | Yes | 77 | 106
187767 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100112 | Error 106 with a 0.9mm Turbo Elite. calibrated fine but as soon as they stepped on the pedal to start lasing, the error would come out.As suggested, Turned the PLS off and back on to reset it, reconnected the 0.9 fiber, calibrated (w/o issues) and error came up again as soon as they stepped on the pedal to lase.Opened a 1.4mm Turbo Elite and it worked as expected.Case completed with laser | FSE found sensor drift and alignment issues, calibrated and aligned, system is operating properly. | No Injury | 20222383 | During Use in Patient | Yes | 84 | 106
187803 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100016 | Error Codes 106, 108 ,201 ,221 ,252 ,310 ,500, 501 & 502 found during the review of the System Log | During a routine service the FSE noted multiple errors in the system log, not duplicated, completed calibrations and system is operating properly. | No Patient Present | 20221972 | Routine Service/Not during a case | Yes | 81 | 106
187865 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Calibration | Failed to maintain calibration | Non-Reportable | Closed | Yes | Tamil Nadu | INDIA (IND) | 100095 | 2.5mm catheter calibration Unsuccessful and 0.9mm catheter is throwing error 106 while lasing | FSE found condensation on the optics and sensors, cleaned and completed calibrations system is operating properly. Instructed customer on storage condition requirements. | No Injury | 20222433 | During Preparation for Use | Yes | 78 | 106
188049 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 240 | Fault 240 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | SWEDEN (SWE) | 100011 | This complaint was opened to document unreported error code 240 blower fault that was discovered during log review during a Preventive Maintenance on 17 August 2022. The error 240 was cleared after a restart and did not affect any treatment nor caused any delays. | During routine service the FSE noted an E240 in the system log, not duplicated. Completed calibrations and the system is operating properly. | No Patient Present | 20220240-BV | Routine Service/Not during a case | Yes | 76 | 240
188189 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100112 | Error Codes 106 & 401 found during the review of the System Log. | During routine service the FSE noted E106 in the system log, not duplicated, completed calibrations and system is operating properly. | No Patient Present | 20222383 | Routine Service/Not during a case | Yes | 76 | 106
188280 | Unknown fault code displayed | Fault 106 | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CO | UNITED STATES (USA) | 100070 | Unknown fault with a 1.4mm catheter. FSE to inform how the case was completed.Received information from FSE via email on 19 Sep 2022:After the laser failed the doctor used another device (I don’t know what product) and treated the patient same day. They continued treating the patient without delay. | FSE found the OA was misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20222413 | During Use in Patient | Yes | 74
188313 | Unknown fault code displayed | Low Voltage Power Supply | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Rep said case was completed with no problems. After the case was completed laser gave a fault which they didn’t remember and upon restarting the laser it did not fully boot up. It stayed on the blue screen with no errors noted. Case was completed with the laser. | FSE found a defective low voltage power supply, replaced. Completed calibrations and system is operating properly. | No Injury | 20222415 | After the Procedure | Yes | 75
188376 |  |  | Fault 230 |  |  |  | Closed | No | AZ | UNITED STATES (USA) | 100059 | Error 230 and 240 found after review of system logs |  | No Patient Present | N/A | Routine Service/Not during a case |  | 0 | 230
188377 | Multiple Fault Codes Displayed | Fault 240 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100059 | Error 230 and 240 found after review of system logs | During service the FSE noted an E240 and 230 in the system logs, not duplicated, completed calibrations and system is operating properly. | No Patient Present | N/A | Routine Service/Not during a case | Yes | 73 | 230
188799 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | SWEDEN (SWE) | 100007 | This is a retroactive report complaint.Logs were downloaded on 08/31/2022 and showed that there was an unreported error code FC 230 on 07/31/2021. | During a routine service the FSE noted an E230 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20212495-BV | Routine Service/Not during a case | Yes | 105 | 230
188810 | Clinical catheter fails to calibrate | Electro / Mechanical | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100083 | Laser would not calibrate, no energy detected on calibration | FSE found a defective CES sensor, replaced. Completed calibrations and system is operating properly. | No Injury | 20222435 | During Preparation for Use | Yes | 83
189086 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100117 | Laser gave a E106 when trying to calibrate. Case completed by other means | FSE was not able to duplicate the E106, did note the coupler fixture failed the switch test and replaced. Completed calibrations and system is operating properly. | No Injury | 20222603 | During Preparation for Use | Yes | 76 | 106
189191 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Hardware | Display / Monitor issue | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100068 | Failed clock battery on the HIM. | During routine service the FSE noted a defective battery in HMI, replaced HMI. Completed calibrations and system is operating properly. | No Patient Present | 20222265 | Routine Service/Not during a case | Yes | 75
189402 |  |  | Fault 500 |  |  |  | Closed | No | AZ | UNITED STATES (USA) | 100083 | Error 500 and 501 found during review of system log. |  | No Patient Present | 20221459 | Routine Service/Not during a case |  | 1 | 500
189403 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100083 | Error 500 and 501 found during review of system log. | During a routine service the FSE noted an E500 and 501 in the service log, not duplicated. Replaced the shutter per the CR, Completed calibrations and system is operating properly. | No Patient Present | 20222435 | Routine Service/Not during a case | Yes | 76 | 500
189627 | Clinical catheter fails to calibrate | Environmental Conditions | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | Tenlangana | INDIA (IND) | 100076 | 2.5mm reference catheter is showing results as Unsuccessful Calibration while trying to calibrate.Multiple attempts were made but every time system shows error Unsuccessful calibration.System was switched ON for checking 1 day before the procedure was scheduled. | FSE found condensation on the sensors and optics, cleaned. Completed calibrations and system is operating properly. | No Patient Present | 20222641 | Not during a case | Yes | 75
189711 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100093 | E108 during the procedure. Reset the laser and completed the case. Informed Sales Rep about E108 and the delay required for the footswitch. | TS determined the E108 is due to depressing the footswitch too quickly after selecting ready button. System is operating properly. | No Injury | N/A | During Preparation for Use | Yes | 72 | 108
189928 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100080 | Error Code 108 found during the review of the System Log. | During routine service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222324 | Routine Service/Not during a case | Yes | 70 | 108
190217 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100008 | On reviewing logs during PM of today, there is one instance of FC 230 on 23 May 2022 that was not reported. Opening compliant in order to document this occurrence. Issue was investigated during PM and no fault found. | During a routine service the FSE noted an E230 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | Non-dispatch | Routine Service/Not during a case | Yes | 65 | 230
190333 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IA | UNITED STATES (USA) | 100097 | Rep called to report an E103 with a 2.3mm catheter, case was completed by other means at the same facility.Postmarket Surveillance reached out to Field Service Manager on 06 Oct 2022 by text to clarify case detail; FS Manager stated that the patient was treated by other means during the same procedure. | FSE verified the E103 and determined the coupling efficiency of the rail was low, cleaned optics, aligned, and completed calibrations. Also replaced customers reference catheter which coupled low. System is operating properly. | No Injury | 20222686 | During Preparation for Use | Yes | 64 | 103
190463 | Clinical catheter fails to calibrate | No Power From Outlet | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes |  | SPAIN (ESP) | 100004 | Laser is not working. Will follow up on this.ADDITIONAL INFORMATION OBTAINED FROM FIELD SERVICE ON 25 OCT 2022:I already made the service on that laser. It was starting properly. It seams that they blowed the fuse when restart the laser. Postmarket Surveillance received additional information via email from Field Service Engineer on 25OCT2022:The patient wasn't even on the table(see file attachment)INFORMATION OBTAINED BY FSE 13 OCT 2022:The patient was treated with our mechanical tool and they noticed the issue during preparation for use. The postponement was a wrong information. Only one patient was affected | FSE contacted distributor and determined the facility power was tripped, system powers on and operates properly. | No Injury | 20220498-BV | During Preparation for Use | Yes | 62
190519 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100128 | This is for install for PLS 100128. Was found bad during install. Per the FSE, PN P016555 (Assembly, PES) was found bad during installation. | During installation the FSE noted a defective PES sensor, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20222665 | During Installation | Yes | 62
191054 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | when using a 1.7 catheter Laser gave a E103. Rep then used a 2.0 and completed the case. Rep stated that E103 has happened before this week but was not called in to report it. | FSE was not able to duplicate the E103, system calibrated and operated properly on arrival. | No Injury | 20222918 | During Preparation for Use | Yes | 61 | 103
191260 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Unable to replicate complaint | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | The original complaint was called in by Rep Larry Vadivia on 10/14 for a E103, complaint # 191054.Rep stated that this happened also on 10/12. When calibrating E103 occurred. He was able to calibrate catheter successfully and complete the case | FSE was not able to duplicate the reported problem, system operated properly on arrival and completed service without errors. | No Injury | 20222918 | During Preparation for Use | Yes | 57 | 103
191282 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | I installed PLS SN 100133 at Pulse Vascular yesterday. Sales Rep Robert Wolcott did the in-service and they are using the laser for cases already. POC Adam Mahler states during case with 0.9 catheter they got error 106 while running the catheter at 50-50. They were able the finish the case with the laser and patient is ok. I’ll go back to the account today to troubleshoot the problem. | FSE verified E106 and adjusted attenuator step with attenuator zero routine, completed calibrations and system is operating properly. | No Injury | 20222923 | During Use in Patient | Yes | 51 | 106
191337 | The laser system covers are damaged | Damaged Laser System Cover | N/A | Hardware | Enclosure - Structure cracks/breaks/gaps | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100123 | Rep noticed a crake in the laser while doing an inservice on this new laser. Laser is working properly. No case involved Contact: Rep Anthony Chevaire @ 603-860-2038 | FSE verified damaged cover, cracked rear handle and side cover replaced both. Completed calibrations and system is operating properly. | No Patient Present | 20222925 | Not during a case | Yes | 56
191549 | Multiple Fault Codes Displayed | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100017 | Error code 500, 501, 211, 230 found during review of system logs | During a routine service the FSE noted several errors in the logs, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222565 | Routine Service/Not during a case | Yes | 47 | 500
191558 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100086 | Error 108 and 252 found after review of system logs | During a routine service the FSE noted error 108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | Non-dispatch | Routine Service/Not during a case | Yes | 53 | 108
191728 | Multiple Fault Codes Displayed | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | Error 108, 211, 210, 221, 401, 252, and 103 found during review of system logs | During a routine service the FSE noted several errors in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222918 | Routine Service/Not during a case | Yes | 43 | 108
191773 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 240 | Fault 240 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | BH7 7DW | UNITED KINGDOM (GBR) | 100002 | During a routine service visit, upon reviewing the logs I noted one instance of an error that was not due to standard operation. The error code was 240 - Blower Fault. There was only one event @14:33 on 20/07/2022.On the next service I will assess if further investigation is required dependent on whether the error re-occurs. | during a routine service the FSE noted an E240 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | Non-dispatch | Routine Service/Not during a case | Yes | 71 | 240
191815 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100098 | The laser faulted out 5 time with the FC 106. Patient was treated by another device. Stents. Competitors device. | FSE completed sensor calibration, optical alignment and system calibrations. System is operating properly. | No Injury | 20222985 | During Use in Patient | Yes | 50 | 106
192314 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100093 | Error Code 108 found during the review of the System Log - 6/9/2022 | During a routine service the FSE noted an E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222581 | Routine Service/Not during a case | Yes | 44 | 108
192326 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100093 | rror Code 108 found during the review of the System Log - 9/15/2022 | During a routine service the FSE noted an E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222581 | Routine Service/Not during a case | Yes | 62 | 108
192329 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100093 | Error Code 108 found during the review of the System Log - 9/22/2022 | During a routine service the FSE noted an E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222581 | Routine Service/Not during a case | Yes | 62 | 108
192330 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100093 | Error Code 108 found during the review of the System Log - 9/26/2022 | During a routine service the FSE noted an E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222581 | Routine Service/Not during a case | Yes | 62 | 108
192341 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100142 | Rep called to report an E106 at 30/25 and 30/30 using a .9 catheter, case was completed at higher settings without issue. | FSE was not able to correct the E106 at the low settings while using a .9 device, system operates properly at all other settings. System will need a software update once released to correct this issue. | No Injury | 20223033 | During Use in Patient | Yes | 62 | 106
192355 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100085 | Error code 106 and 302 found during review of system logs | FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222782 | Routine Service/Not during a case | Yes | 77 | 106
192623 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Co | UNITED STATES (USA) | 100075 | Error 106 found during review of system logs | FSE noted an E106 in the system logs, found beam waste was too small, increased size. Completed calibrations and system is operating properly. | No Patient Present | 20222258 | Routine Service/Not during a case | Yes | 57 | 106
192666 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100000 | During the planned Preventive Maintenance Service, the FSE discovered that the Calibration Energy Sensor was functioning intermitently. It was determined that it is needed to be replaced. The parts were ordered and will be replaced on delivery to Leusden BV. | During a routine service the FSE noted a defective CES sensor, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20220234-BV | Routine Service/Not during a case | Yes | 56
192697 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100139 | Error 106 during installation | During installation the FSE noted the VES and PES sensors had a delta greater than 10%, replaced both. Completed calibrations and system is operating properly. | No Patient Present | 20222942 | During Installation | Yes | 56 | 106
192795 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100159 | E106 when trying to use the laser. Case completed with drug coated balloon | FSE was not able to duplicate the reported problem. Completed calibrations and system is operating properly. | No Injury | 20223079 | During Preparation for Use | Yes | 74 | 106
192911 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Hardware | Display / Monitor issue | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100132 | When performing an in-service Rep had a very hard time adjusting fluence on touchscreen. | FSE verified touchscreen not responsive, replaced HMI. Completed calibrations and system is operating properly. | No Patient Present | 20223083 | Not during a case | Yes | 56
192933 | Multiple Fault Codes Displayed | Fault 106 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100090 | Error 105, 106, 108, 240, 252, and 310 found after review of system logs | FSE noted several errors in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222488 | Routine Service/Not during a case | Yes | 56 | 310
192936 |  |  | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | AZ | UNITED STATES (USA) | 100147 | Error 106 found during review of system logs |  | No Patient Present | 20222975 | During Installation |  | 29 | 106
192948 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100024 | Error Codes 500, 501 found during the review of the System Log. | FSE noted an E500 in the system log during service, replaced shutter per CR. Completed calibrations and system is operating properly. | No Patient Present | 20223010 | Routine Service/Not during a case | Yes | 56 | 500
193066 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100150 | Seth said they got a 106 fault code during the case, reset laser and tried again and got the same fault. Case was completed with Balloon , patient is doing well. | System has multiple 106 errors and FSE was not able to rectify, system will be replaced. | No Injury | 20223087 | During Use in Patient | Yes | 58 | 106
193112 | The laser system covers are damaged | Rare Occurrence | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100065 | Lauren said the door on the laser is not closing properly and the customer would like to have someone to look at it. Case was completed with laser and is doing well. | FSE found hardware too tight, loosened bolts and catheter coupler door operated properly. Completed calibrations and system is operating properly. | No Injury | 20223089 | During Preparation for Use | Yes | 64
193194 | Multiple Fault Codes Displayed | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100080 | Jeff (TM) said they got a 500, 501 200 and 201 error codes when they turned on the laser. Case was completed with rotorblade and balloon. Patient is doing well. | FSE found a loose connection to the attenuator motor, reconnected. Completed calibrations and system is operating properly. | No Injury | 20223094 | During Preparation for Use | Yes | 63 | 500
193206 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100053 | Error 108 found during routine service | FSE noted an E108 in the system log during routine service, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222463 | Routine Service/Not during a case | Yes | 63 | 108
193280 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100159 | Arlen said the account got a 106 error code while using the laser. Case was completed with their second laser. Patient is doing well. | Customer refused service, was using the system without issue. FSE closed the work order. | No Injury | 20223104 | During Use in Patient | Yes | 64 | 106
193342 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Other | Other | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100148 | Difference between service detector and VES > 10% | During installation the FSE found the VES detector >10% delta, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20222992 | During Installation | Yes | 64
193421 | The laser system indicates fault codes 400-499 - Heartbeat (HB) faults (NEX) | Fault 401 | Fault 401 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100150 | Reviewing uploaded log from complaint # 193066 found error code 401 happened on 11/9 | Non dispatch during routine service the FSE noted an E401 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223087 | Not during a case | Yes | 63 | 401
193422 |  |  | Fault 250 |  |  |  | Closed | No | MD | UNITED STATES (USA) | 100150 | Reviewing log that was uploaded for complaint #193066 error code 250 happened on 11/11 |  | No Patient Present | N/A | Not during a case |  | 74 | 250
193426 |  |  | Fault 250 |  |  |  | Closed | No | MD | UNITED STATES (USA) | 100150 | Review of log uploaded for complaint 193066 an error code 250 was found on 11/14 |  | No Patient Present | N/A | Not during a case |  | 74 | 250
193435 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100042 | Reviewing log uploaded from PM code 500 and 501 happened on 5/9 | Non dispatch, during routine service the FSE noted an E500 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20221422 | Not during a case | Yes | 63 | 500
193436 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 501 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100042 | Reviewing log on recent PM found errors 500 and 501 happened on 6/10 | Non dispatch, during routine service the FSE noted a E500 in the logs, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20221422 | Not during a case | Yes | 63 | 501
193440 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100042 | Reviewing log that was uploaded for PM error 108 happened on 9/26 | Non dispatch, during a routine service the FSE noted an E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20221422 | Not during a case | Yes | 63 | 108
193450 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100150 | Laser gave E106 when trying to calibrate | FSE was not able to duplicate the reported problem. Completed calibrations and system is operating properly. | No Injury | 20223274 | During Preparation for Use | Yes | 63 | 106
193455 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100051 | Reviewing log that was uploaded during PM found 108 happened on 9/30 | During a routine service the FSE noted an E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222880 | Not during a case | Yes | 63 | 108
193457 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CT | UNITED STATES (USA) | 100074 | Diana said they got a 106 Error on their laser , they completed the case with angioplasty. patient is doing well. | FSE found the alignment out, realigned. Completed calibrations and system is operating properly. | No Injury | 20223277 | During Use in Patient | Yes | 63 | 106
193458 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100051 | Reviewing uploaded log from PM a 108 error was recorded on 10/27 | During a routine service the FSE noted an E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222880 | Not during a case | Yes | 63 | 108
193464 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100156 | Sinan said the account got a 103 Error. | FSE was not able to duplicate the reported problem, system calibrated and operated properly. | No Injury | 20223278 | During Preparation for Use | Yes | 63 | 103
193576 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100150 | Seth said they got a 106 fault code again on their PLS laser. Case was completed with Balloon and patient is doing well. | FSE was not able to rectify the E106, system was deinstalled and replaced with SN 100174 | No Injury | 20223282 | During Use in Patient | Yes | 62 | 106
193809 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100132 | E106 during the procedure. Case was completed with their CVX300. | FSE worked with TS to determine the VES and PES exhibit noisy signals, replaced both. Completed calibrations and system is operating properly. | No Injury | 20223292 | During Use in Patient | Yes | 57 | 106
194351 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 211 | Fault 211 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100026 | Customer called to report an exception 211 during the procedure. Case was completed using the CVX laser on site. | FSE found a defective shutter, replaced. Completed calibrations and system is operating properly. | No Injury | 20223305 | During Use in Patient | Yes | 82 | 211
194468 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100104 | Error Code 108 found during the review of the System Log. | During routine service the FSE noted an E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222874 | Routine Service/Not during a case | Yes | 52 | 108
194566 | The laser system will not turn on | Keyswitch | Unknown | Power | Will not turn on | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100059 | No error code. The laser would not turn on. | FSE found a defective keyswitch, replaced. Completed calibrations and system is operating properly. | No Injury | 20223327 | During Preparation for Use | Yes | 51
194690 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100157 | Error: 106This happened during a case.Case was not completed by laser. They used another laser from the same facility to complete the case (from IR). | FSE verified fault and adjusted attenuator step position, completed calibrations and system is operating properly. | No Injury | 20223334 | During Use in Patient | Yes | 50 | 106
194747 | The laser system turns off during a case | Fault 106 | Unknown | Power | Other | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100118 | System would shut down when they stepped on the pedal. | FSE verified fault and adjusted the attenuator step position, completed calibrations and system is operating properly. | No Injury | 20223335 | During Preparation for Use | Yes | 49
194848 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100089 | Will not calibrate with .9... tried multiple .9 catheters from different lots with no success | FSE found optical rail required alignment, aligned and completed calibrations. System is operating properly. | No Injury | 20223340 | During Preparation for Use | Yes | 57
194905 |  |  | N/A |  |  |  | Closed | No | MA | UNITED STATES (USA) | 100148 | While performing installation procedure, found VES failed sensor cal at 11%. Replaced VES sensor...complete install successfully |  | No Patient Present | 20222992 | During Installation |  | 0
195090 | Multiple Fault Codes Displayed | Fault 106 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100059 | Error 106 and 108 found during review of system logs | FSE noted E106 and 108 in the system logs during routine service, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223327 | Not during a case | Yes | 53 | 108
195094 |  |  | Fault 103 |  |  |  | Closed | No | TX | UNITED STATES (USA) | 100094 | Error codes 103, 108 and 252 that were found in the system |  | No Patient Present | 20222861 | Routine Service/Not during a case |  | 3 | 103
195095 | Multiple Fault Codes Displayed | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100094 | Error codes 103, 108 and 252 that were found in the system | During a routine service the FSE noted Error 103 and 108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222861 | Routine Service/Not during a case | Yes | 53 | 103
195286 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100091 | Jennifer said the laser gave them a fault code 106 and they were unable to get past it. Completed case with Tightrail and patient is doing well. | FSE found the attenuator zero step off, adjusted. Completed calibrations and added isolators per CR, system is operating properly. | No Injury | 20223367 | During Use in Patient | Yes | 51 | 106
195305 | Multiple Fault Codes Displayed | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100026 | Error codes found after review of system log 221 Atten PLS Fault, 500 Self test fail, 501 Power up timeout | During service the FSE noted an E221, 500, 501. Not duplicated during service, completed calibrations and system is operating properly. | No Patient Present | 20223305 | Routine Service/Not during a case | Yes | 57 | 221
195347 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 500 | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100010 | Dear all,Attached you can find a Complaint for PLS - S/N 100010 - installed in Azienda Ospedaliera di Sassari – Italy On reviewing of logs, there was 1 instance of FC 211 on 16 Sep 2022. This occurred during a customer activity not reported.I've checked the system fully function. The error occurred for shutter assembly that will be replaced when the part will be in stock.I have opened a complaint in order to document that occurrence.Kind Regards, | During routine service the fSE noted E500 and 211, not duplicated however parts were not available. Will return and replace shutter per CR. Completed calibrations and system is operating properly. | No Patient Present | N/A | Routine Service/Not during a case | Yes | 56
195393 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Savannah said the laser gave a Fault code 106 during the case, reset laser and tried another turbo elite catheter and was able to complete the case. patient is doing well.Note: FSIS shows this account name as Tenet Fl Physician Services, the name has changed to Prime Vascular Institute the address above is correct. | FSE verified 106, completed isolator installation per CR, completed calibrations and system is operating properly. | No Injury | 20223549 | During Use in Patient | Yes | 56 | 106
195491 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100117 | Kevin said they got a Fault 106 on their laser during the case , Completed case with a competor's laser (Angio Dynamics). Patient is doing well. | FSE adjusted attenuator zero step count and added isolators per CR, completed calibrations and system is operating properly. | No Injury | 20223556 | During Use in Patient | Yes | 55 | 106
195578 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100154 | David: During a case today, the PLS at NYP Columbia suffered a fault code 108 after calibration when transitioning to ready mode. The machine was restarted and recalibrated and the case was completed successfully. Patient is doing well | FSE did not duplicate the reported problem, completed calibrations and system is operating properly. | No Injury | 20223563 | During Use in Patient | Yes | 52 | 108
195588 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100159 | Turbo Power failed during procedure 106 error code. They turned the laser on and off and re-calibrated with no success, they got another Turbo Power and completed the procedure. | Customer refused service, states system is operating properly and continues to use it without further issue. | No Injury | 20223615 | During Use in Patient | Yes | 52 | 106
195603 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CO | UNITED STATES (USA) | 100102 | During routine service noted E500 in the system log. Not duplicated, completed CR 156370, shutter replacement. | During routine service the FSE noted an E500 in the log, completed shutter replacement per CR. Completed calibrations and system is operating properly. | No Patient Present | N/A | Routine Service/Not during a case | Yes | 52 | 500
195604 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CO | UNITED STATES (USA) | 100102 | During a routine service an E108 was noted in the log, not duplicated. | During a routine service the FSE noted an E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20222849 | Routine Service/Not during a case | Yes | 52 | 108
195848 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | Error Code 108 found during the review of the System Log on 6/22/2022. | During routine service the FSE noted E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223138 | Routine Service/Not during a case | Yes | 56 | 108
195851 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | Error Code 108found during the review of the System Log on 6/29/2022. | During routine service the FSE noted E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223138 | Routine Service/Not during a case | Yes | 56 | 108
195853 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | Error Code 108 found during the review of the System Log on 7/6/2022. | During routine service the FSE noted E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223138 | Routine Service/Not during a case | Yes | 56 | 108
195856 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | Error Codes 103 & 108 found during the review of the System Log on 8/10/2022. | During routine service the FSE noted E103 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223138 | Routine Service/Not during a case | Yes | 56 | 103
195858 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | Error Code 103 found during the review of the System Log on 8/17/2022. | During routine service the FSE noted E103 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223138 | Routine Service/Not during a case | Yes | 56 | 103
195860 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | Error Code 108 found during the review of the System Log on 9/19/2022. | During routine service the FSE noted E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223138 | Routine Service/Not during a case | Yes | 56 | 108
195863 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | Error Code 108 found during the review of the System Log on 11/23/2022. | During routine service the FSE noted E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223138 | Routine Service/Not during a case | Yes | 56 | 108
196175 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Loose Hardware | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100089 | Rep called to report an E500 and E501, case was completed by other means at the same time and facility. | FSE found PCA's not seated completely, seated. Completed calibrations and system is operating properly. | No Injury | 20223625 | During Preparation for Use | Yes | 50 | 500
196218 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100015 | Error Codes 108 & 303 found during the review of the System Log on 12/28/2022. | During routine service the FSE noted E108 and E303 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223165 | Routine Service/Not during a case | Yes | 49 | 108
196245 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100190 | Rep called to report an E106 with 2 different .9 catheters, calibrated fine, tried several different fluence and rate settings but gave a 106 every time. Case was completed with balloon at the same time and facility. | FSE did not duplicate the E106, completed calibrations and system is operating properly. | No Injury | 209223644 | During Preparation for Use | No | 48 | 106
196307 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Code 106 found during the review of the System Log on 12/7/2022 16:58 and 17:04 | During service the FSE noted additional E106 in the system logs, completed CR162895 to add isolators to the VES/PES sensors. Completed calibrations and system is operating properly. | No Patient Present | 20223549 | Not during a case | Yes | 43 | 106
196308 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 105 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Code 105 found during the review of the System Log on 11/2/2022 12:17 | During service the FSE noted an E105 in the system log, not duplicated. completed calibrations and system is operating properly. | No Patient Present | 20223549 | Not during a case | Yes | 43 | 105
196309 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Code 108 found during the review of the System Log on 5/2/22 5:16, 5/25/22 13:32, 6/27/22 13:05, 9/19/22 13:51, 10/24/22 15:44, 11/4/22 14:24, 12/12/22 17:26, 5/2/22 13:06 | During service the FSE noted an E108 in the system log, not duplicated. completed calibrations and system is operating properly. | No Patient Present | 20223549 | Not during a case | Yes | 43 | 108
196310 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 201 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Code 201 found during the review of the System Log on 5/9/22 12:30 | During service the FSE noted an E201 in the system log, not duplicated. completed calibrations and system is operating properly. | No Patient Present | 20223549 | Not during a case | Yes | 43 | 201
196311 | Multiple Fault Codes Displayed | Fault 221 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Codes 201, 221, & 310 found during the review of the System Log on 6/16/2022 | During service the FSE noted several errors in the system log, not duplicated. completed calibrations and system is operating properly. | No Patient Present | 20223549 | Not during a case | Yes | 52 | 201
196312 | Multiple Fault Codes Displayed | Fault 201 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Codes 201, 221 & 240 found during the review of the System Log on 6/28/2022 | During service the FSE noted an E201, 221, 240 in the system log, not duplicated. completed calibrations and system is operating properly. | No Patient Present | 20223549 | Not during a case | Yes | 52 | 201
196313 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 303 | Fault 303 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Code 303 found during the review of the System Log on 7/13/22 12:47 | During service the FSE noted an E303 in the system log, not duplicated. completed calibrations and system is operating properly. | No Patient Present | 20223549 | Not during a case | Yes | 52 | 303
196320 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100149 | FC E106 occured when trying to use the laser. Laser was reset and error cleared. Case completed with laser. | FSE was not able to duplicate the reported problem, completed update per CR162895, isolator install. Completed calibrations and system is operating properly. | No Injury | 20230009 | During Preparation for Use | Yes | 52 | 106
196344 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100089 | Rep called in a error 500 and 501. Case completed by other means | FSE found a defective attenuator assembly, replaced. completed calibrations and system is operating properly. | No Injury | 20230005 | During Preparation for Use | Yes | 52 | 500
196385 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IA | UNITED STATES (USA) | 100097 | Laser gives an E103 when using a 2.3 turbo elite catheter. Case completed with laser. | FSE found coupling low, cleaned focus lens and adjusted alignment. Completed calibrations and system is operating properly. | No Injury | 20230011 | During Use in Patient | Yes | 56 | 103
196442 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100103 | Error Code 108 found during the review of the System Log | During service the FSE noted an E108 in the system log, not duplicated. completed calibrations and system is operating properly. | No Patient Present | 20222799 | Routine Service/Not during a case | Yes | 55 | 108
196447 | Other - Justification Required | Manufacturing Error | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100187 | Laser Keys and Interlock Plug not found for installation. | During installation the FSE noted the keys and interlock connector missing, ordered and completed installation. | No Patient Present | 20223571 | During Installation | Yes | 84
196465 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100186 | Laser gave a E103 when using a 2.0 turbo elite catheter | FSE was not able to duplicate the error but noted the alignment was adjusted. Completed calibration and system is operating properly. | No Injury | 20230023 | During Use in Patient | Yes | 55 | 103
196517 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100034 | Reportable log errors for complaints.Fault 200 12/27/22 15:28250. 12/27/22. 16:33106. 12/30/22. 14:13 12/30/22. 14:18, 14:25303. 12/30/22. 14:13 | During service the FSE noted an E106 and E303 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223186 | Routine Service/Not during a case | Yes | 55 | 106
196545 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100161 | When using the laser a E108 occurred. Case was completed with laser | FSE was not able to duplicate the E108 while onsite for complaint 196602. Completed calibrations and system is operating properly. | No Injury | 20230031 | During Use in Patient | Yes | 54 | 108
196602 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100157 | Jennifer said they got 106 fault , opened a new catheter and recalibrated and got the same 106 code. Case was completed with Balloon. Patient is doing well. | FSE was not able to duplicate the E106 on arrival but did get a 106 after system set up, adjusted attenuator offset. Completed calibrations and system is operating properly. | No Injury | 20230040 | During Preparation for Use | Yes | 67 | 106
196690 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100132 | Kelly said they got a 106 code, so they reset laser and recalibrated the catheter and was able to complete the case with no other issues. Patient is doing well.Kelly Kershaw ph# 757-472-6154 | FSE completed alignments and system calibrations, system is operating properly. | No Injury | 20230046 | During Use in Patient | Yes | 65 | 106
196697 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MN | UNITED STATES (USA) | 100039 | NL 93 requiring complaint 108 | During routine service the FSE noted an E108 in the logs, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223396 | Routine Service/Not during a case | Yes | 65 | 108
196760 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100190 | William (TM) said they got a 106 code on their laser , reset laser and tried several different catheters before getting one to work, Case was completed with laser. Patient is doing well.William Brumfield @ 719-650-4039 Call When you have service Scheduled. | FSE verified 106, system requires a new PES but will be deinstalled due to customer confidence and replaced with SN 100205 | No Injury | 20230049 | During Use in Patient | Yes | 64 | 106
196826 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100047 | Error found during review of system log10/28/2022 16:41 NL93 EXCEPTION [230] | During routine service the FSE noted an E230 in the logs, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223479 | Routine Service/Not during a case | Yes | 64 | 230
196827 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100044 | Error found during review of system log8/29/2022 15:38 NL93 EXCEPTION [252] | During a routine service the FSE noted an E252 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223478 | Routine Service/Not during a case | Yes | 64 | 252
196829 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IA | UNITED STATES (USA) | 100097 | NL 93 e103 found after review of system logs | During the corrective maintenance the FSE noted a 103 in the system log. Completed calibrations and the system is operating properly. | No Patient Present | 20230011 | Routine Service/Not during a case | Yes | 64 | 103
196889 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100197 | E106 with a 2.0mm catheter. Sales Rep restarted the laser with the same result. Case was completed with their CVX300. This was a peripheral case. | FSE verified the 106, adjusted alignment and attenoffset, completed calibrations and system is operating properly. | No Injury | 20230063 | During Use in Patient | Yes | 63 | 106
196997 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 231 | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100055 | Error Codes 231 & 301 found during the review of the System Log | During a routine service the FSE noted error 231 and 301 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223403 | Routine Service/Not during a case | Yes | 62 | 231
197015 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100100 | E108 during the case. Laser restarted and catheter recalibrated with no issue. Case completed with the laser. | Non dispatch, TS instructed customer about E108, system is operating properly. | No Injury | N/A | During Use in Patient | Yes | 60 | 108
197038 | The laser system has broken components | Electro / Mechanical | N/A | Hardware | Mechanical Movement | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | During inspection, it was found that one of the front casters will not unlock. This did not affect any cases. | FSE verified damaged wheel, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20230077 | Not during a case | Yes | 61
197228 | The laser system indicates fault codes 300-399 - State faults (NEX) | Unable to replicate complaint | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CO | UNITED STATES (USA) | 100070 | E310 during the procedure. Laser was restarted and the case was completed with the laser. | FSE was not able to duplicate the reported problem. Completed calibrations and system is operating properly. | No Injury | 20230263 | During Use in Patient | Yes | 59 | 310
197438 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | Rep called to report an E106 with a 14 FR Glidelight, case was completed at the same time with the CVX laser on site. | FSE verified 106, adjusted atten offset, completed calibrations and system is operating properly. | No Injury | 20230271 | During Use in Patient | Yes | 57 | 106
197455 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MI | UNITED STATES (USA) | 100120 | Mark called to report an E108, reset the system and completed the case without issue. | Non dispatch, TS informed customer concerning E108, system is operating properly. | No Injury | N/A | During Use in Patient | Yes | 57 | 108
197463 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Catheter Coupler Assy | Fault 200 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100155 | Tieger reported a E200 and 201 with a 1.4mm catheter, reference catheter and .9 operates properly. Training laser used for demos at Philips San Diego. No patients are involved. | FSE found a defective A34 in coupler assembly, removed pin 1, not used. Training laser, A34 will be replaced on next visit. Completed calibrations and system is operating properly. | No Patient Present | 20230274 | Not during a case | Yes | 68 | 200
197479 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100018 | The hospital send this request to our general Philips email address and they are asking about the fault 106. after further investigation from Josef Tholemeyer-Noring the OR staff said that the fault came when they are starting the laser or maybe by calibration. The laser is due for an installation as the evaluation agreement is over now and the system is with the biomed department now and we need to arrange the pickup, sending an empty crate and the deinstallation procedureCS note: System does not have a service contract and the systemis going to be de-installed. | System will be deinstalled as evaluation agreement is over, system will be repaired at Philips. | No Patient Present | 2023002-BV | During Preparation for Use | Yes | 56 | 106
197483 |  |  | Fault 250 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | TN | UNITED STATES (USA) | 100046 | Error Code 250 found during the review of the System Log on 8/30/2022 |  | No Patient Present | 20223255 | Routine Service/Not during a case |  | 24 | 250
197484 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100046 | Error Code 252 found during the review of the System Log on 8/15/2022 | During routine service the FSE noted a E252 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223255 | Routine Service/Not during a case | Yes | 59 | 252
197536 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | SC | UNITED STATES (USA) | 100192 | E106 with a 0.9mm catheter. Laser was restarted and the case completed with the laser. | FSE verified and adjusted the atten offset one step, completed calibrations and system is operating properly. | No Injury | 20230275 | During Use in Patient | Yes | 56 | 106
197650 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100202 | Jodi said the account got a 106 Fault on their PLS laser . Case was completed with Balloon. Patient is doing well. | FSE verified and replaced VES and PES sensors, completed calibrations and system is operating properly. | No Injury | 20230303 | During Use in Patient | Yes | 65 | 106
197652 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100197 | Error Code 106 found during the review of the System Log on 12/23/22 @ 12:10, 12:15 | During service the FSE noted E106 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230063 | Routine Service/Not during a case | Yes | 53 | 106
197655 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100197 | Error Code 106 found during the review of the System Log on 12/23/22 @ 18:46, 18:48, 18:51. | During service the FSE noted E106 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230063 | Routine Service/Not during a case | Yes | 53 | 106
197656 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100197 | Error Code 310 found during the review of the System Log on 1/13/23 @ 17:44. | During service the FSE noted E310 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230063 | Routine Service/Not during a case | Yes | 53 | 310
197671 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | HI | UNITED STATES (USA) | 100027 | fault code 500 and a fault code 501 found during review of system log | During routine service the FSE noted a E500 and 501 in the system log, date of data pull suggests this was during the service visit 6/28-6/30. Completed calibrations and system is operating properly. | No Patient Present | 20223157 | Routine Service/Not during a case | Yes | 53 | 500
197704 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100185 | Don called to report an E106, case was completed with the CVX on site. This was a peripheral case. | FSE verified and adjusted the atten offset, completed calibrations and system is operating properly. | No Injury | 20230280 | During Use in Patient | Yes | 56 | 106
197748 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | SC | UNITED STATES (USA) | 100192 | Error Code 106 found during the review of the System Log on 1-19-23. | During service the FSE noted E106 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230275 | Routine Service/Not during a case | Yes | 56 | 106
197750 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | SC | UNITED STATES (USA) | 100192 | Error Code 106 found during the review of the System Log on 1-18-23. | During service the FSE noted E106 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230275 | Routine Service/Not during a case | Yes | 65 | 106
197802 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100132 | Error Codes 106, 240 & 310 found during the review of the System Log on 12/03/22. | During service the FSE noted E106, 240, 310 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230046 | Routine Service/Not during a case | Yes | 64 | 106
197807 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100132 | Error Codes 106, 221, 505, 310 & 302 found during the review of the System Log on 12/7/22. | During service the FSE noted E106 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230046 | Routine Service/Not during a case | Yes | 64 | 106
197809 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100132 | Error Code 106found during the review of the System Log on 1/10/23. | During service the FSE noted E106 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230046 | Routine Service/Not during a case | Yes | 64 | 106
197812 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100132 | Error Code 106 found during the review of the System Log on 1/11/23. | During service the FSE noted E106 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230046 | Routine Service/Not during a case | Yes | 64 | 106
197815 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100132 | Error Code 221 found during the review of the System Log on 12/4/22. | During service the FSE noted E221 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20230046 | Routine Service/Not during a case | Yes | 64 | 221
197936 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | E106 when laser was turned on | FSE verified and adjusted the atten offset, completed calibrations and system is operating properly. | No Injury | 2023088 | During Preparation for Use | Yes | 69 | 106
198012 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log on 7/8/2022 2:08:04 PM | During service the FSE noted E108 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20223528 | Routine Service/Not during a case | Yes | 62 | 108
198014 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log on 7/22/2022 3:25:40 PM | During service the FSE noted E108 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20223528 | Routine Service/Not during a case | Yes | 62 | 108
198016 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log on 11/10/2022 5:46:52 PM | During service the FSE noted E108 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20223528 | Routine Service/Not during a case | Yes | 62 | 108
198019 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log on 1/13/2023 2:24:45 PM | During service the FSE noted E108 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20223528 | Routine Service/Not during a case | Yes | 62 | 108
198022 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 240 | Fault 240 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 240 found during the review of the System Log on 1/20/2023 3:43:59 PM | During service the FSE noted E240 in the system log, completed calibrations and system is operating properly. | No Patient Present | 20223528 | Routine Service/Not during a case | Yes | 60 | 240
198055 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100202 | The catheter was calibrated as normal but failed to perform past 10 seconds in the body. The system read error code 106. The laser catheter would no longer perform after the inital 10 seconds and the error message. | Customer has lost confidence and system will be replaced with SN 100218. | No Injury | 20230303 | During Use in Patient | Yes | 80 | 106
198137 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100205 | During installation the FSE was getting a 103 and required replacement of the PES to correct. | During installation the FSE noted a noisy PES sensor, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20230257 | Routine Service/Not during a case | Yes | 59 | 103
198144 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IL | UNITED STATES (USA) | 100054 | NL93 requiring complaints: 108 found after review of system log | During a routine service the FSE noted a 108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230026 | Routine Service/Not during a case | Yes | 58 | 108
198145 | The laser system will not turn on | Interlocks Intermittently Dis-Engage | N/A | Power | Will not turn on | Non-Reportable | Closed | Yes | IA | UNITED STATES (USA) | 100097 | No case. Customer heard buzzing sound when they came in Monday morning. No Patient involved | FSE verified rear interlock tripped, adjusted rear panel. Completed calibrations and system is operating properly. | No Patient Present | 20230344 | Not during a case | Yes | 85
198207 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | E106 when trying to use the laser. Case completed with 2nd laser on site | FSE verified and replaced the attenuator, completed calibrations and system is operating properly. | No Injury | 20230345 | During Preparation for Use | Yes | 64 | 106
198219 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100093 | While using on a patient at settings 65/45 a E106 error occurred. Settings were lowered and case was completed with laser | FSE verified, installed isolators per CR, adjusted alignment completed calibrations, system is operating properly. | No Injury | 20230346 | During Use in Patient | Yes | 64 | 106
198264 | The laser system generates unusual noise | Interlocks Intermittently Dis-Engage | N/A | Other | Other | Non-Reportable | Closed | Yes | CO | UNITED STATES (USA) | 100167 | When laser is plugged in an intermittent buzzing sound occurred. Laser was moved to different out let and buziing stopped. Laser was used for case | FSE found rear interlock disengaged, adjusted rear cover. Completed calibrations and system is operating properly. | No Injury | 20230356 | During Preparation for Use | Yes | 63
198347 | The laser system generates unusual noise | Interlocks Intermittently Dis-Engage | N/A | Other | Other | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100086 | laser 100086 was buzzing when the customer came into the lab in the morning. The only way they could get it to stop was to turn the breaker off or un plug the laser. There were no cases or patients involved. | FSE found rear interlock dis engaged, replaced the rear panel. Completed calibrations and system is operating properly. | No Patient Present | 20230358 | Not during a case | Yes | 62
198381 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100042 | Error Code 310 found during the review of the System Log on 10/13/2022 @ 23:42 | During service the FSE noted a 310 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230330 | Routine Service/Not during a case | Yes | 70 | 310
198409 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100026 | E106 when trying to use the laser | FSE was not able to duplicate, added isolators per CR and completed calibrations. System is operating properly. | No Injury | 20230361 | During Preparation for Use | Yes | 70 | 106
198465 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100213 | During installation got error 106. Mike and Tim helped me resolve the problem by replacing the PES. | During installation the FSE found a noisy PES sensor giving E106, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20230286 | During Installation | Yes | 69 | 106
198535 | Multiple Fault Codes Displayed | Fault 201 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Florida | UNITED STATES (USA) | 100185 | Error Code 201 found during the review of the System Log at 1603Error code 221 found during the review of the System Log 1553 and again at 1558 | During a service the FSE noted an error 201 and 221 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230280 | Routine Service/Not during a case | Yes | 69 | 201
198588 | The laser system does not provide reliable performance | Electro / Mechanical | Unknown | Power | Will not turn on | Non-Reportable | Closed | Yes |  | FRANCE (FRA) | 100066 | during the preparation of the case , technician has constated a crash of the software of PLS .Calibration has been made with the laser prob test and after with the Elca 0,9 for the procedure. She has noticed a blue screen that has staid longer on the philips logoShe has turned off one time and she has called me to fixe the problem.We have made several operations but the system doesn't work at all. | FSE found a defective ECM PCB, replaced. Completed calibrations and system is operating properly. | No Injury | 20230048-BV | During Use in Patient | Yes | 66
198682 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Tenlangana | INDIA (IND) | 100076 | Error 106 while starting the clinical procedure. Customer tried recalibrating catheter and continue with procedure but system continue to throw error 106. Customer couldn't use LASER for the clincal procedure and completed procedure without LASER. | FSE found dust particles and sticky substance on the Optics and sensors, cleaned. Added isolators per CR and completed calibrations, system is operating properly. | No Injury | 20230383 | During Use in Patient | Yes | 66 | 106
198768 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100148 | Mike said they got a 106 code during the case , reset laser and recalibrated was able to complete the case with laser but they got the 106 code a second time at the end. Patient is doing well. | FSE verified, added isolators per CR and adjusted atten offset, completed calibrations and system is operating properly. | No Injury | 20230385 | During Use in Patient | Yes | 65 | 106
198992 | Clinical catheter fails to calibrate | Fault 106 | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100071 | Ken (FSE) said PLS laser would not calibrate after moving the laser to Cath Lab 2 to prep for a case.TM Lisa Saylor tried different plugs and cleaning the detectorTried calibrating my reference catheters and the CVX reference catheter.Unable to use for the case. Doctor choose to use a mechanical device from our competitors to complete the case. patient is doing well. | FSE was unable to calibrate a catheter, found E106 in the logs also. Completed calibrations and system is operating properly. | No Injury | 20230401 | During Preparation for Use | Yes | 78
199014 | Multiple Fault Codes Displayed | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100093 | Error Code 103 found during the review of the System Log at 1648Error Code 302 found during the review of the System Log at 1648 | During service the FSE noted errors 103 and 302 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230346 | Routine Service/Not during a case | Yes | 63 | 103
199018 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 1000093 | Error Code 106 found during the review of the System Log on 1/30/2023 | During service the FSE noted error 106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230346 | Routine Service/Not during a case | Yes | 63 | 106
199029 | Multiple Fault Codes Displayed | Fault 221 | Fault 200 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100093 | Error Code 200 found during the review of the System Log at 1701Error Code 221 found during the review of the System Log at 1705Error Code 310 found during the review of the System Log at 1702Error Code 401 found during the review of the System Log at 1653 | During service the FSE noted errors 200, 221, 310, 401 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230346 | Routine Service/Not during a case | Yes | 63 | 200
199032 |  |  | Fault 200 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | NC | UNITED STATES (USA) | 100093 | Error Code 200 found during the review of the System Log on 11/1/2022 @ 1259 |  | No Patient Present | 20230346 | Routine Service/Not during a case |  | 22 | 200
199038 |  |  | Fault 200 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | NC | UNITED STATES (USA) | 100093 | Error Code 200 found during the review of the System Log at 1259 |  | No Patient Present | 20230346 | Routine Service/Not during a case |  | 22 | 200
199046 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100071 | Error Code 106 found during the review of the System Log on 1/25/23 @ 3:47pm | FSE noted an E106 in the system log not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230401 | Routine Service/Not during a case | Yes | 63 | 106
199047 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100071 | Error Code 106 found during the review of the System Log on 2/7/23 @ 12:20pm | FSE noted E106 in the system log not duplicated, completed calibrations and system is operating properly. | No Patient Present | 20230401 | Routine Service/Not during a case | Yes | 63 | 106
199089 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 303 | Fault 303 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | SPAIN (ESP) | 100004 | On reviewing logs during PM of today, there is one instance of FC 303 of an “Invalid State_ UI” on 9th January 2023 that was not reported. Opening compliant in order to document this occurrence. | During a routine service the FSE noted a E303 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | N/A | Routine Service/Not during a case | Yes | 74 | 303
199133 | Clinical catheter fails to calibrate | Environmental Conditions | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | Tamil Nadu | INDIA (IND) | 100095 | 2.5mm catheter calibration Unsuccessful while testing PLS system. | FSE found contamination on optics and sensors, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20230404 | Not during a case | Yes | 62
199160 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | 12/2/2022 18:11 NL93 EXCEPTION [252] found during review of system log | During service the FSE noted an E252 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230127 | Routine Service/Not during a case | Yes | 62 | 252
199162 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | 12/23/2022 14:29 NL93 EXCEPTION [106] found during review of system log. | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230127 | Routine Service/Not during a case | Yes | 62 | 252
199163 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | 12/20/2022 22:40 NL93 EXCEPTION [252] found during review of system log | During service the FSE noted an E252 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230127 | Routine Service/Not during a case | Yes | 62 | 252
199164 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | 1/16/2023 14:29 NL93 EXCEPTION [106] found during review of system log. | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230127 | Routine Service/Not during a case | Yes | 62 | 106
199166 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | 1/16/2023 14:29 NL93 EXCEPTION [106] 1/19/2023 16:55 NL93 EXCEPTION [106]1/19/2023 16:55 NL93 EXCEPTION [201]1/19/2023 16:57 NL93 EXCEPTION [106]1/19/2023 17:37 NL93 EXCEPTION [106]found during review of system log. | During service the FSE noted an E106 and E201 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230127 | Routine Service/Not during a case | Yes | 62 | 106
199221 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | SPAIN (ESP) | 100004 | On reviewing logs during PM of today, there is one instance of FC 106 of an “Sensor Mismatch” on 2nd December 2022 that was not reported. Opening compliant in order to document this occurrence. | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230071-BV | Routine Service/Not during a case | Yes | 59 | 106
199240 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100218 | E106 with a Turbo Power 2.0mm catheter. Laser was reset and the case was completed with the laser at the same hospital on the same day. | FSE noted a noisy PES sensor and replaced. Completed calibrations and system is operating properly. | No Injury | 20230413 | During Use in Patient | Yes | 59 | 106
199318 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | CZECH REPUBLIC (CZE) | 100082 | Error 106. | FSE did not duplicate the reported error but attached system log verifying E106, completed alignment and calibrations, system is operating properly. | No Patient Present | 20230106-BV | Not during a case | Yes | 58 | 106
199443 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100016 | Error Code 106 found during the review of the System Log on 09/28/22 @ 12:33 | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230140 | Routine Service/Not during a case | Yes | 57 | 106
199445 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100098 | E106 during the case. Case was not completed with the laser, it was completed with an atherectomy device on the same day at the same facility. No harm to patient. | FSE verified E106, adjusted alignment and added isolators per CR to the sensor brackets. Completed calibrations and system is operating properly. | No Injury | 20230574 | During Use in Patient | Yes | 57 | 106
199446 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100016 | Error Code 108 found during the review of the System Log on 09/07/22 @ 18:34 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230140 | Routine Service/Not during a case | Yes | 57 | 108
199451 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100016 | Error Code 108 found during the review of the System Log on 09/30/22 @ 17:50 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230140 | Routine Service/Not during a case | Yes | 57 | 108
199453 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100016 | Error Code 108 found during the review of the System Log on 10/27/22 @ 15:07 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230140 | Routine Service/Not during a case | Yes | 69 | 108
199454 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100016 | Error Code 230 found during the review of the System Log on 02/09/23@ 17:12 | During service the FSE noted an E230 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230140 | Routine Service/Not during a case | Yes | 69 | 230
199455 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100016 | Error Code 310 found during the review of the System Log on 01/11/23 on 16:49 | During service the FSE noted an E310 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230140 | Routine Service/Not during a case | Yes | 69 | 310
199636 | Clinical catheter fails to calibrate | Bad Clinical Catheter | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100141 | Failure to calibrate on a 0.9mm catheter. Second 0.9mm catheter calibrated successfully and the case was completed with the laser, same day and facility. | FSE found damaged clinical catheter, completed calibrations and system is operating properly with FSE catheters. | No Injury | 20230583 | During Preparation for Use | Yes | 75
199886 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100051 | Rep called to report an exception 103, case was completed with balloon. | FSE was not able to duplicate the reported problem, completed calibrations and system is operating properly. | No Injury | 20230604 | During Preparation for Use | Yes | 63 | 103
199933 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100071 | Rep called to report a E106, system was rebooted and calibrated. Was able to get one pass, case was completed by other means at the same time and facility. | FSE verified a 106, adjusted the attenuator offset, completed calibrations and system is operating properly. | No Injury | 20230606 | During Use in Patient | Yes | 70 | 106
200026 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100205 | REP stated:I need to report a PLS system failure that was experienced yesterday at Northwest Medical Center (2801 FL-7, Margate, FL 33063). This is a replacement PLS, as the first model also experienced the Error 106 System Failure code. QTY 3 of 0.9 laser fibers were used to try and complete the case and troubleshoot the system, nothing worked. The physician could not complete the case with laser atherectomy and ultimately ended up bailing on the procedure. The (multiple) physicians that requested this upgraded laser console are very unhappy with their experience—they have not had a single case that the laser has actually worked during, and that’s not just in the IR lab for PV, that’s also the physicians utilizing laser in the coronary procedures. This is unacceptable as they are threatening to request removal of the system and replace with Auryon laser.Case was not completed with the laser and patient will be rescheduled at a later date, no complications were reported. | FSE verified E106, noted noise on PES and adjusted alignment. Completed calibrations and system is operating properly. | No Injury | 20230613 | During Use in Patient | Yes | 61 | 106
200141 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100185 | Don called to report an error 103, system was recalibrated and case was completed without issues. | FSE found alignment out slightly, aligned. Completed calibrations and system is operating properly. | No Injury | 20230620 | During Use in Patient | Yes | 60 | 103
200153 | The laser system has broken components | Electro / Mechanical | N/A | Hardware | Mechanical Movement | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100233 | Caster(s) broken on arrival. | FSE noted a damaged front caster during removal from crate for installation, replaced. Completed calibration and system is operating properly. | No Patient Present | 20230581 | Routine Service/Not during a case | Yes | 60
200175 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100098 | Customer reported a 106, case was completed by other means at the same time and facility. | FSE verified the 106, adjusted alignment and completed calibrations. System is operating properly. | No Injury | 20230621 | During Preparation for Use | Yes | 60
200279 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100205 | Error Code 103 found during the review of the System Log on 1/29/2023 @ 0:46, 0:48 & 0:51. | During service the FSE noted a E103 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230613 | Routine Service/Not during a case | Yes | 57 | 103
200292 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100041 | Error Code E103 occurred 11/28/22 found during the review of the System Log at 14:07 | During service the FSE noted a E103 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20223385 | Routine Service/Not during a case | Yes | 57 | 103
200377 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100045 | Error Code 106 found during the review of the System Log on 09/28/22 @ 12:23 | During service the FSE noted a E106 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230152 | Routine Service/Not during a case | Yes | 59 | 106
200379 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100045 | Error Code 108 found during the review of the System Log on 09/07/22 @ 18:34 | During service the FSE noted a E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230152 | Routine Service/Not during a case | Yes | 59 | 108
200380 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100045 | Error Code 108 found during the review of the System Log on 09/30/22 on 17:50 | During service the FSE noted a E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230152 | Routine Service/Not during a case | Yes | 59 | 108
200382 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100045 | Error Code 108 found during the review of the System Log on 10/27/22 at 15:07 | During service the FSE noted a E108 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230152 | Routine Service/Not during a case | Yes | 59 | 108
200383 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100045 | Error Code 230 found during the review of the System Log on 02/08/23 @ 17:12 | During service the FSE noted a E230 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230152 | Routine Service/Not during a case | Yes | 59 | 230
200384 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100045 | Error Code 310 found during the review of the System Log on 01/11/23 @ 16:49 | During service the FSE noted a E310 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230152 | Routine Service/Not during a case | Yes | 59 | 310
200396 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100051 | E106 when trying to use the laser. Case completed with balloon | FSE verified and installed isolators per CR. Completed calibrations and system is operating properly. | No Injury | 20230631 | During Preparation for Use | Yes | 72 | 106
200487 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | WA | UNITED STATES (USA) | 100119 | E106 when trying to calibrate. Case completed with 2nd laser on site | FSE was not able to duplicate the reported problem but noted a bad reference catheter, replaced. Completed calibrations and system is operating properly. | No Injury | 20230646 | During Preparation for Use | Yes | 58 | 106
200533 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100213 | E106 when using 1.7 catheter at highest settings. Recalibrated and completed case using settings at 50/70. | FSE was not able to duplicate, completed calibrations and system is operating properly. | No Injury | 20230643 | During Use in Patient | Yes | 57 | 106
200556 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100185 | When calibrating laser E103 occurred. Laser was recalibrated and E302 happened. Case completed with 2nd laser on site | FSE found rail out of alignment, aligned. Completed calibrations and system is operating properly. | No Injury | 20230647 | During Preparation for Use | Yes | 57 | 103
200884 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MN | UNITED STATES (USA) | 100039 | Brent said the account got a 106 error code , they were able to reset the laser and recalibrate and complete the case. Patient is doing well. | FSE found noisy sensors, completed isolator update per CR. Completed calibrations and system is operating properly. | No Injury | 20230679 | During Use in Patient | Yes | 52 | 106
201114 | Clinical catheter fails to calibrate | Electro / Mechanical | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100225 | Leanne said she went to do some training on the newly installed laser and she keeps getting a Calibraion Unsuccessful message and can't get past it. Leanne McDaniels (TM) 540-435-1311 . She wants the FSE call her with a time to service their first cases are scheduled for sometime next week.Thanks | Upon arrival, the FSE found that the laser was not firing. FSE found that the fiber optic connector on the PCN was not fully seated. FSE reseated the connector and completed the PVI with no further issues. System operating properly at the conclusion of the PVI. | No Patient Present | 20230705 | Not during a case | Yes | 50
201176 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Tamil Nadu | INDIA (IND) | 100095 | Error 106 while using 0.9mm catheter for clinical use.When Customer noticed error, they rebooted system and retried and found LASER is firing without error. but reconfirm customer LASER is ready, Customer rebooted for 2nd time and after calibration when they tried firing LASER system throws error 106, Customer tried rebooting system for 5 attempts but got not succeed as error 106 continues with 0.9mm catheter.Customer replaced catheter to check but even with 2nd catheter of 0.9mm, system popes error 106. | Upon arrival, the FSE confirmed the E106 condition. During inspection of the laser, the FSE found contamination on the PES sensor. FSE cleaned the surface of the PES and completed the PVI and multiple E106 tests with no issue. System meets all Philips specifications. | No Injury | 20230693 | During Use in Patient | Yes | 63 | 106
201294 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100023 | Matt said they were getting an Unsuccessful calibration with the .9 clinical catheter so they changed to the 1.4 clinical catheter and were able to complete the case. patient is doing well. | FSE confirmed that a 0.9mm catheter failed calibration. FSE also found that there are scuff marks on the side of the covers that can account for the optical rail misalignment. FSE completed optical rail alignment and PVI with no issues and tested 0.9mm catheters multiple times with no failed calibrations. System is operating properly. | No Injury | 20230696 | During Preparation for Use | Yes | 49
201374 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IL | UNITED STATES (USA) | 100226 | E106 and E303 with a 0.9 catheter. Case was not completed with the laser, it was completed by other means at the same hospital without delay. | FSE adjusted the spot size and attenuator offset position. Completed calibrations and system is operating properly. | No Injury | 20230707 | During Use in Patient | Yes | 46 | 106
201436 | Multiple Fault Codes Displayed | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100185 | E103 during the procedure followed by E301 and E106. Case was not completed with the PLS. It was completed with their CVX300 on the same day at the same facility without delay. | System will be deinstalled and replaced with SN 100250. | No Injury | 20230710 | During Use in Patient | Yes | 45 | 103
201462 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Shutter Assembly | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | E500/501 at startup, would not clear with multiple restarts. Case not completed with the laser. It was completed with a Phoenix device, same day at same facility with no delay. | FSE verified E500 and replaced the defective shutter. Completed calibrations and system is operating properly. | No Injury | 20230711 | During Preparation for Use | Yes | 45 | 500
201465 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100248 | Problem: coupler board is missing a contact arm. | During installation the FSE found a damaged coupler, A34 switch, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20230668 | During Installation | Yes | 45
201469 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes |  | CZECH REPUBLIC (CZE) | 100082 | The PLS failed to calibrate the sterile catheter they were about to use on a patient.The customer could not calibrate their non sterile reference catheter when they switched on the system. They could, however, calibrate their sterile catheter, which they used to successfully treat the patient with the PLS system. They were able to use the PLS system, SN100082. | FSE found the optical rail mis aligned, all optic adjustments were tight. Front cover has some scratches and was likely mishandled. Completed calibrations and system is operating properly. | No Injury | 20230156-BV | During Preparation for Use | Yes | 45
201479 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Optical Rail Mis-Aligned | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100098 | EC103 & EC 200 during an In-Service training session. No patient involvement. | FSE found optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Patient Present | 20230715 | Not during a case | Yes | 73 | 103
201914 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | UNITED KINGDOM (GBR) | 100204 | During the installation of the Laser an Error 500 was experienced. On consultation with the Tier 3 support it was agreed that even though a new style shutter assembly was installed it was the most likely culprit of the error. This will be replaced ASAP | FSE found a defective shutter assembly, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20230145-BV | During Installation | Yes | 44 | 500
201943 | Multiple Fault Codes Displayed | Fault 221 | Fault 401 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MN | UNITED STATES (USA) | 100039 | FSE reported E401, 221, 201 in the system log dated 05 January 2023, found during service on 13 March. | During service the FSE noted E221, 201, and 401 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230679 | Not during a case | Yes | 44 | 401
201960 | Multiple Fault Codes Displayed | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100023 | Error Code 108 found during the review of the System Log @ 1:10:23 PMError Code 500 found during the review of the System Log @ 1:11:21 PMError Code 501 found during the review of the System Log @ 1:11:46 PM | During service the FSE noted E108, 500, 501 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230696 | Routine Service/Not during a case | Yes | 44 | 108
201969 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100189 | Rep called to report an E230, rebooted the system but the error would not clear. Case was completed at the same time and place by other means. | FSE found J5 connection loose, connected. Completed calibrations and system is operating properly. | No Injury | 20230887 | During Preparation for Use | Yes | 44 | 230
202226 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | E106 with a 2.0mm catheter. Sales Rep reset laser, recalibrated the same catheter and completed the case with the laser without delay. | Reported as a 106 but FSE noted error was a 108, not duplicated. Completed calibrations and system is operating properly. | No Injury | 20230908 | During Use in Patient | Yes | 41 | 106
202243 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100233 | Error 108, case was not completed with the laser. It was completed with angio and balloon without delay at the same facility. | Non dispatch phone fix, customer depressing footswitch within 2 seconds of pressing ready. System is operational. | No Injury | N/A | During Use in Patient | Yes | 41 | 108
202257 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100202 | 1-20-23, Error Code 105, VES out of range | During service the FSE noted an E105 in the system log, not duplicated. The system will be deinstalled due to multiple 106 issues and replaced by SN 100218. | No Patient Present | N/A | Routine Service/Not during a case | Yes | 39 | 106
202258 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100090 | 11-16-22 Error 106 Sensor Mismatch error found during review of system logs | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | N/A | Routine Service/Not during a case | Yes | 39 | 106
202259 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100090 | 01-17-23 Error 108 Max Invalid pulses error found during review of system logs | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230043 | Routine Service/Not during a case | Yes | 39 | 108
202323 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | DENMARK (DNK) | 100166 | During a training session the laser was failing with the FC106. No patient present | FSE completed the CR for updating with sensor isolators, completed calibrations and system is operating properly. | No Patient Present | 20230073-BV | Not during a case | Yes | 38 | 106
202325 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100200 | E106. Laser was reset and recalibrated the catheter. Case was completed with the laser without delay. | FSE was not able to duplicate the reported problem, verified E106 in system log, completed isolator upgrade and calibrations and system is operating properly. | No Injury | 20230936 | During Use in Patient | Yes | 38 | 106
202769 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100113 | E106 when trying to calibrate. Case completed with Balloon | FSE did not duplicate the reported problem, noted adjusting spot size and alignment. Completed calibrations and system is operating properly. | No Injury | 20230997 | During Preparation for Use | Yes | 34 | 106
202866 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Unable to replicate complaint | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Rob said the laser would not calibrate their catheters getting a 106 code and unsuccessful calibration error. Case will be completed with their other laser. Patient is doing well.Account contact is Lisa McGovern 631-534-7246. Rob Wilcott TM 412-860-2258 Rob said they are very busy lab and it maybe hard to reach Lisa , FSE can come an work on laser as they can pull it out of the room and let them work on it. | FSE was not able to duplicate the reported problem. Completed calibrations and system is operating properly. | No Injury | 20230969 | During Preparation for Use | Yes | 31 | 106
203127 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100230 | Jonathan said his sales rep texted him that the laser at Gomelb had a fault 106 yesterday. The account was able to restart and recalibrate and complete the case with the laser with no more faults. | FSE found a noisy PES sensor, replaced, and adjusted the attenoffset position. Completed calibrations and system is operating properly. | No Injury | 20230989 | During Preparation for Use | Yes | 29 | 106
203137 | The laser system does not provide reliable performance | Did Not Know How To Operate Laser | N/A | Other | Other | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100225 | Kevin said -Laser red aiming light is not on. laser is operating properly and case was completed with laser | FSE indicated the system is operating properly, customer did not know the system must be in ready for the aiming diode to turn on. | No Injury | 20230993 | During Preparation for Use | Yes | 29
203272 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100113 | Account said they got a Fault codes 310/350. Completed lead removal case with mechanical device. Patient is doing ok. | FSE found a defective footswitch, replaced. Also noted a damaged 45 mirror mount, replaced. completed calibrations and system is operating properly. | No Injury | 20230997 | During Preparation for Use | Yes | 27 | 310
203398 | The laser system has malfunctioning components | Low Max Energy < 3 Months | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Rep called to report a message that no energy is detected during calibration. The case was completed with the CVX laser system on site. | FSE found low max energy, system was deinstalled and replaced with SN 100264. | No Injury | 20231006 | During Preparation for Use | Yes | 37 | 103
203449 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100250 | Rep called to report a E600 during a procedure, switched to the CVX to complete the case. | FSE found wiring ties pulling on fuse holder, cut the tie, completed calibrations and system is operating properly. | No Injury | 20231015 | During Use in Patient | Yes | 36 | 600
203469 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100169 | Rep called to report an error 103 with a .9 and 2.0 device. Case was completed by other means and the same facility and time. | FSE verified E106, adjusted the LVPS and completed calibrations, system is operating properly. | No Injury | 20231017 | During Use in Patient | Yes | 36 | 103
203502 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MI | UNITED STATES (USA) | 100108 | Rep contacted FSE to report the system would not calibrate the 2.5 reference catheter. Clinical catheter was pulled and calibrated and case was completed without issue. | FSE verified failed calibrations, adjusted optical rail alignment and found marginal reference catheter, replaced. Completed calibrations and system is operating properly. | No Injury | 20231041 | During Preparation for Use | Yes | 23
203661 | Multiple Fault Codes Displayed | Fault 240 | Fault 240 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | Error Code 240 found during the review of the System Log – 2/21/23 – 12:37Error Code 252 found during the review of the System Log – 2/21/23 – 13:31 | During service the FSE noted an E240 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230711 | Not during a case | Yes | 21 | 240
203664 | Multiple Fault Codes Displayed | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100105 | Error Code 252 found during the review of the System Log – 3/6/2023 – 12:44Error Code 105 found during the review of the System Log – 3/6/2023 – 14:48 | During service the FSE noted an E252 and E105 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230711 | Not during a case | Yes | 22 | 252
203667 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 105 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 105 found during the review of the System Log – 11/21/2022 – 15:08 | During service the FSE noted an E105 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 22 | 105
203668 | Multiple Fault Codes Displayed | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Error Code 103 found during the review of the System Log – 10/3/2022 – 14:50Error Code 108 found during the review of the System Log – 10/3/2022 – 12:59 | During service the FSE noted an E103 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230544 | Routine Service/Not during a case | Yes | 22 | 103
203669 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 105 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 105 found during the review of the System Log – 1/12/2023 15:29 | During service the FSE noted an E105 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | N/A | Not during a case | Yes | 22 | 105
203670 | Multiple Fault Codes Displayed | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Error Code 103 found during the review of the System Log – 2/3/23 – 15:48Error Code 108 found during the review of the System Log – 2/3/23 – 19:12 | During service the FSE noted an E103 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230544 | Routine Service/Not during a case | Yes | 22 | 103
203672 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 105 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 105 found during the review of the System Log – 1/30/2023 15:53 | During service the FSE noted an E105 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 22 | 105
203673 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Error Code 103 found during the review of the System Log – 3/30/23 – 12:56 | During service the FSE noted an E103 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230544 | Routine Service/Not during a case | Yes | 22 | 103
203675 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Error Code 108 found during the review of the System Log – 9/16/22 – 19:19 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230544 | Routine Service/Not during a case | Yes | 22 | 108
203676 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 105 found during the review of the System Log – 2/16/2023 15:46 | During service the FSE noted an E105 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 21 | 108
203678 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Error Code 108 found during the review of the System Log – 11/17/22 – 22:00 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230544 | Routine Service/Not during a case | Yes | 22 | 108
203679 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 105 found during the review of the System Log – 10/28/2022 20:28 | During service the FSE noted an E105 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 22 | 108
203680 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Error Code 108 found during the review of the System Log – 1/9/23 – 19:53 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230544 | Routine Service/Not during a case | Yes | 22 | 108
203681 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 108 found during the review of the System Log 11/15/2022 17:50 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 22 | 108
203682 | Multiple Fault Codes Displayed | Fault 201 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100062 | Error Code 201 found during the review of the System Log – 9/7/22 – 15:14Error Code 201 found during the review of the System Log – 9/7/22 – 15:15Error Code 310 found during the review of the System Log – 9/7/22 – 16:63 | During service the FSE noted an E201 and 310 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230544 | Routine Service/Not during a case | Yes | 25 | 201
203683 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 108 found during the review of the System Log – 11/16/2022 15:58 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 25 | 108
203684 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 105 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 105 found during the review of the System Log – 11/21/2022 18:21 | During service the FSE noted an E105 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 28 | 105
203685 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 105 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 105 found during the review of the System Log – 12/14/2022 15:44 | During service the FSE noted an E105 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 28 | 105
203686 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 108 found during the review of the System Log – 1/4/2023 14:46 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 28 | 108
203687 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 108 found during the review of the System Log – 1/20/2023 16:24 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 28 | 108
203699 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 108 found during the review of the System Log - 2/7/2023 15:53Error Code 108 found during the review of the System Log2/7/2023 16:48 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 22 | 108
203700 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 108 found during the review of the System Log – 2/16/2023 15:44 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 26 | 108
203703 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 108 found during the review of the System Log – 2/28/2023 18:43 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 26 | 108
203704 | Multiple Fault Codes Displayed | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 108 found during the review of the System Log - 3/16/2023 15:42Error Code 252 found during the review of the System Log - 3/16/2023 18:16 | During service the FSE noted an E108 and 252 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 26 | 108
203708 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 108 found during the review of the System Log - 3/27/2023 13:08 | During service the FSE noted an E108 in the system log, no duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 28 | 108
203709 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 230 found during the review of the System Log – 11/7/2022 20:11 | During service the FSE noted an E230 in the log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 26 | 230
203714 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 252 found during the review of the System Log – 11/9/2022 15:26 | During service the FSE noted an E252 in the system log, no duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 26 | 252
203715 | Multiple Fault Codes Displayed | Fault 303 | Fault 303 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100133 | Error Code 303 found during the review of the System Log – 3/31/2023 16:11Error Code 310 found during the review of the System Log – 3/31/2023 16:11 | During service the FSE noted an E303 in the system log, no duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230969 | Not during a case | Yes | 28 | 303
203717 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100200 | Error Code 108 found during the review of the System Log – 2/8/23 – 18:18 | During service the FSE noted an E108 in the system log, no duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230936 | Not during a case | Yes | 28 | 108
203720 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100200 | Error Code 106 found during the review of the System Log – 3/27/23 – 13:55 | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230396 | Not during a case | Yes | 26 | 106
203803 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100085 | Error Code 252 found during the review of the System Log – 2-9-23 | During service the FSE noted an E252 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230751 | Not during a case | Yes | 34 | 252
203871 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | E500 during startup self test. Laser was moved to multiple rooms and reset with the same result. Case was not completed with the PLS, it was completed with their CVX without delay at the same facility. | FSE was not able to duplicate the reported problem, completed calibrations and system is operating properly. | No Injury | 20231229 | During Preparation for Use | Yes | 32 | 500
203887 | The footswitch connector/receptacle is damaged | Unable to replicate complaint | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100157 | The front footswitch connector is misaligned and the footswitch cannot be connected. They used the rear footswitch connector and the case was completed with the laser without delay and at the same facility. | FSE was not able to duplicate the reported problem, receptacle is aligned and working properly. Completed calibrations and system is operating properly. | No Injury | 20231234 | During Preparation for Use | Yes | 32
203902 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 210 | Fault 210 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100017 | “Error Code 210 found during the review of the System Log 10/13/2022 17:51 | During service the FSE noted an E210 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230806 | Not during a case | Yes | 32 | 210
203903 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100017 | Error Code 108 found during the review of the System Log – 1/25/2023 21:01 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230806 | Not during a case | Yes | 32 | 108
203904 | Multiple Fault Codes Displayed | Fault 231 | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100017 | Error Code 231 found during the review of the System Log – 3/8/2023 19:45 Error Code 230 found during the review of the System Log – 3/8/2023 19:45 | During service the FSE noted an E230 and 231 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230806 | Not during a case | Yes | 32 | 231
204001 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 505 | Fault 505 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Missouri | UNITED STATES (USA) | 100137 | Error Code 505 found during the review of the System Log 10/19/2022 17:57 | During service the FSE noted an E505 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230686 | Not during a case | Yes | 31 | 505
204002 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Missouri | UNITED STATES (USA) | 100137 | Error Code 106 found during the review of the System Log – 11/30/2022 14:55 | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230686 | Not during a case | Yes | 31 | 106
204003 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Missouri | UNITED STATES (USA) | 100137 | Error Code 106 found during the review of the System Log – 12/7/2022 20:32 | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230686 | Not during a case | Yes | 31 | 106
204004 | Multiple Fault Codes Displayed | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Missouri | UNITED STATES (USA) | 100137 | Error Code 108 found during the review of the System Log – 12/10/2022 0:42 Error Code 106 found during the review of the System Log – 12/10/2022 0:45 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230686 | Not during a case | Yes | 31 | 108
204005 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Missouri | UNITED STATES (USA) | 100137 | Error Code 106 found during the review of the System Log – 12/20/2022 11:31 | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230686 | Not during a case | Yes | 31 | 106
204006 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Missouri | UNITED STATES (USA) | 100137 | Error Code 106 found during the review of the System Log – 1/6/2023 19:42 | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230686 | Not during a case | Yes | 31 | 106
204008 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Missouri | UNITED STATES (USA) | 100137 | Error Code 106 found during the review of the System Log – 1/17/2023 12:44 | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230686 | Not during a case | Yes | 35 | 106
204009 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Missouri | UNITED STATES (USA) | 100137 | Error Code 103 found during the review of the System Log – 3/3/2023 23:14 | During service the FSE noted an E103 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230686 | Not during a case | Yes | 35 | 103
204016 | Multiple Fault Codes Displayed | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Code 500 found during the review of the System Log – 4/17/2023 12:29Error Code 500 found during the review of the System Log – 4/17/2023 12:30Error Code 501 found during the review of the System Log – 4/17/2023 12:30Error Code 500 found during the review of the System Log – 4/17/2023 12:39Error Code 501 found during the review of the System Log – 4/17/2023 12:39 | During service the FSE noted an E500 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20231229 | Not during a case | Yes | 35 | 500
204017 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Error Code 103 found during the review of the System Log – 4/5/2023 14:08 | During service the FSE noted an E103 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20231229 | Not during a case | Yes | 35 | 103
204051 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100169 | E106 with a 2.0mm catheter. Case wa completed by other means without delay at the same facility. | FSE verified 106 and completed calibrations and system is operating properly. | No Injury | 20231255 | During Use in Patient | Yes | 34 | 106
204053 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100169 | E106 and E3102 with a 0.9mm catheter. Case was not completed with the laser. It was completed by other means without delay at the same facility. The doctor has informed the sales rep that he wants the laser removed from the facility. | FSE verified E106, completed calibrations and system is operating properly. | No Injury | 20231255 | During Use in Patient | Yes | 34 | 106
204154 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 240 | Fault 240 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | SPAIN (ESP) | 100021 | On reviewing logs during PM of today, there is one instance of FC 240 of a "Blower Motor Fault" on 30th November 2022 that was not reported. Opening compliant in order to document this occurrence. | During service the FSE noted an E240 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | N/A | Not during a case | Yes | 53 | 240
204205 | Multiple Fault Codes Displayed | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100125 | Error Code 108 found during the review of the System Log – 2/9/2023 6:56:29 PM Error Code 303 found during the review of the System Log – 2/9/2023 6:58:19 PM | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230829 | Not during a case | Yes | 35 | 108
204206 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 201 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100125 | Error Code 201 found during the review of the System Log – 4/20/2023 1:55:37 PM | During service the FSE noted an E201 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230829 | Not during a case | Yes | 42 | 201
204272 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | Error Code 108 found during the review of the System Log 12/8/2022 14:56 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230807 | Not during a case | Yes | 41 | 108
204274 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | Error Code 108 found during the review of the System Log – 12/16/2022 13:19 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230807 | Not during a case | Yes | 41 | 108
204275 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | Error Code 103 found during the review of the System Log – 1/9/2023 15:36 | During service the FSE noted an E103 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230807 | Not during a case | Yes | 41 | 103
204276 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | Error Code 310 found during the review of the System Log – 2/10/2023 18:18 | During service the FSE noted an E310 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230807 | Not during a case | Yes | 41 | 310
204278 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100067 | Error Code 310found during the review of the System Log – 3/31/2023 11:35 | During service the FSE noted an E310 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230807 | Not during a case | Yes | 41 | 310
204288 |  |  | Fault 301 |  |  |  | Closed | No | TN | UNITED STATES (USA) | 100036 | FC 301/302 during a case. Facility restarted the laser to clear the FC's. Case was completed with the laser without delay at the same facility. |  | No Injury | N/A | During Use in Patient |  | 6 | 301
204289 |  |  | Fault 301 |  |  |  | Closed | No | TN | UNITED STATES (USA) | 100036 | FC 301/302 during a case. Facility restarted the laser to clear the FC's. Case was completed with the laser without delay at the same facility. |  | No Injury | N/A | During Use in Patient |  | 6 | 301
204290 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 301 | Fault 301 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | FC 301/302 during a case. Facility restarted the laser to clear the FC's. Case was completed with the laser without delay at the same facility. | FSE was not able to duplicate the reported problem. Completed calibrations and system is operating properly. | No Injury | 20231281 | During Use in Patient | Yes | 41 | 301
204376 | The reference catheter fails to calibrate | Unknown Fault | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | Michigan | UNITED STATES (USA) | 100108 | Laser will not calibrate even with new reference catheter | System was deinstalled and replaced with SN 100286. | No Injury | 20231309 | During Preparation for Use | Yes | 51
204393 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log – 2/17/2023 16:47 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20231281 | Not during a case | Yes | 38 | 108
204394 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100036 | Error Code 108 found during the review of the System Log – 3/31/2023 15:27 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20231281 | Not during a case | Yes | 38 | 108
204636 | Clinical catheter fails to calibrate | Bad Clinical Catheter | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | N/A | SPAIN (ESP) | 100004 | he laser was failing during calibrating the 0.9 cath. No faultcode. This was during preparation for use and the physician was using the 1.4 ELCA cath. Patient is well. The distributor was missing to notify us earlier. The system should be checked with a 0.9 veryfication cath for troubleshooting. There is a possibility that the calibration was not done properly, the cath was defective or dirt on the cathUpdate: 30Jan2023 The device used was a Turbo Elite not an ELCA. | Complaint opened in error to document a disposable issue. Later classified as a system issue and documented by complaint #199221, CRN 20230071-BV. | No Injury | 20230255-BV | During Preparation for Use | Yes | 51
204637 |  |  | Unknown |  |  |  | Closed | No | N/A | SPAIN (ESP) | 100004 | he laser was failing during calibrating the 0.9 cath. No faultcode. This was during preparation for use and the physician was using the 1.4 ELCA cath. Patient is well. The distributor was missing to notify us earlier. The system should be checked with a 0.9 veryfication cath for troubleshooting. There is a possibility that the calibration was not done properly, the cath was defective or dirt on the cath |  | No Injury | 20230255-BV | During Preparation for Use |  | 1
204646 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Michigan | UNITED STATES (USA) | 100108 | E106 | FSE not able to duplicate, system will be deinstalled and replaced due to customer confidence. | No Injury | 20231309 | During Preparation for Use | Yes | 48 | 106
204917 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IL | UNITED STATES (USA) | 100257 | Bob Hedberg - I got a call from the Rep, Michelle in Illinois, Thursday. Laser was getting 106 errors on the 1.7mm catheter at both 45 and 50 fluence. Patient was treated with Balloon and stent and is doing well. | FSE found a noisy PES detector, replaced. Completed calibrations and system is operating properly. | No Injury | 20231338 | During Use in Patient | Yes | 43 | 106
205136 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100104 | Error Code 103 found during the review of the System Log – 12/19/2022 18:02 | During service the FSE noted an E103 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20231192 | Not during a case | Yes | 35 | 103
205137 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100104 | Error Code 106 found during the review of the System Log – 12/6/2022 - 14:07 | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20231192 | Not during a case | Yes | 35 | 106
205140 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100104 | Error Code 106 found during the review of the System Log – 2/9/2023; 13:45 | During service the FSE noted an E106 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20231192 | Not during a case | Yes | 35 | 106
205141 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100104 | Error Code 108 found during the review of the System Log – 1/26/2023; 14:21 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20231192 | Not during a case | Yes | 35 | 108
205142 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100104 | Error Code 108 found during the review of the System Log – 3/29/2023; 17:22 | During service the FSE noted an E108 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20231192 | Not during a case | Yes | 35 | 108
205347 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Telangana | INDIA (IND) | 100231 | Fault Code:106 towards the end of the Procedure with 0.9mm catheter.User retried calibrating 0.9mm after error but Calibration resulted Unsuccessful Calibration. after 2 tries, Calibration passed but after couple of cycles during procedure again system Lasing stopped with error 106.Customer was performing 1st case on the LASER system after its installation. | FSE verified E106, completed alignment, calibration and attenuator zero routines, system is operating properly. | No Injury | 20231408 | During Use in Patient | Yes | 44 | 106
205505 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Fault 106 | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100248 | Rep called to report an E600 post procedure. Case was completed without issue. | FSE found problem with wire routing to the DC isolator, re routed/straightened. Completed calibrations and system is operating properly. | No Injury | 20231423 | After the Procedure | Yes | 34 | 600
205526 | Clinical catheter fails to calibrate | Electro / Mechanical | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100194 | Rep called to report the PLS indicated no energy detected during calibration. Case was cancelled and will be rescheduled. | FSE found PES not reading data, reseated the PCBs, completed calibrations and system is operating properly. | No Injury | 20231429 | During Preparation for Use | Yes | 34
205601 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Hardware | Display / Monitor issue | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100249 | Rep Rob was doing an inservice and the techs notified him that they have been having issues with the lasers diplay buttons / touchscreen.No case or patient involment. | FSE found a defective HMI, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20231432 | Not during a case | Yes | 41
205752 |  |  | Fault 108 |  |  |  | Closed | No | TX | UNITED STATES (USA) | 100237 | Greg Norman site Philips Rep called to report that the site tried to use the PLS Laser System but the system gave a fault during calibration. They were not sure what the fault was but thought it might have been a fault 108.They decided not to use the PLS Laser and used the CVX300 to do the case with.The patient did well with the CVX300 Laser System. |  | No Injury | N/A | During Preparation for Use |  | 2 | 108
205753 |  |  | Fault 108 |  |  |  | Closed | No | TX | UNITED STATES (USA) | 100237 | Greg Norman site Philips Rep called to report that the site tried to use the PLS Laser System but the system gave a fault during calibration. They were not sure what the fault was but thought it might have been a fault 108.They decided not to use the PLS Laser and used the CVX300 to do the case with.The patient did well with the CVX300 Laser System. |  | No Injury | N/A | During Preparation for Use |  | 2 | 108
205754 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Optical Rail Mis-Aligned | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100237 | Greg Norman site Philips Rep called to report that the site tried to use the PLS Laser System but the system gave a fault during calibration. They were not sure what the fault was but thought it might have been a fault 108.They decided not to use the PLS Laser and used the CVX300 to do the case with.The patient did well with the CVX300 Laser System. | FSE found optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20231567 | During Preparation for Use | Yes | 38 | 108
205867 | The power cord cable is damaged | Damaged Power Cord | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100199 | Philips Rep William Follensbee reported that they were unable to do the case because the PLS laser would shut off if the power cord was moved. They tried reconnecting it at the laser and the wall outlet but it kept shutting off if the cord was moved. Site decided to use another device for the case.The Philips Rep is requesting that we send overnight early delivery a new power cord the PLS System.Send to the Cath Lab contact: Jason Daniels 706-280-7705AC Power Cord NEMA 5-20P Locking 15ft Part # P015805 | Non dispatch, replacement power cord was sent to the customer. Did not fix the issue, see complaint #205983 for dispatch repair. | No Injury | 20231570 | During Preparation for Use | Yes | 51
205896 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100283 | Field Sales Rep Matt Tulgan reported that the site Tech Cynthia Guzman reported that the system gave a fault 106 and a 303 while trying to calibrate a catheter. They decided to use another device to do the case.Site Contact: Cynthia Guzman 619-488-6048 | FSE adjusted alignment and attenuator offset, completed calibrations and system is operating properly. | No Injury | 20231572 | During Preparation for Use | Yes | 51 | 106
205982 |  |  | N/A |  |  |  | Closed | No | GA | UNITED STATES (USA) | 100199 | The site Philips Rep William Follensbee requested that we over night a new PLS Power Cord to the site.The site received the new power cord but they still have the same problem. The system shuts off shortly after it is turned on. They hear a humming noise coming from the system even though it is off.The site is requesting a FSE to look at the system. |  | No Injury | 20231569 | During Preparation for Use |  | 14
205983 | The laser system has malfunctioning components | Damaged AC Inlet | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100199 | The site Philips Rep William Follensbee requested that we over night a new PLS Power Cord to the site.The site received the new power cord but they still have the same problem. The system shuts off shortly after it is turned on. They hear a humming noise coming from the system even though it is off.The site is requesting a FSE to look at the system. | FSE found a damaged AC inlet receptacle, replaced. Completed calibrations and system is operating properly. | No Injury | 20231569 | During Preparation for Use | Yes | 51
206143 | Multiple Fault Codes Displayed | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100104 | The system started giving fault 103's and 303's during the case.They stopped using the laser and completed the case with a balloon.The patient was doing fine after the case. | FSE was not able to duplicate the reported issue, adjusted attenuator offset and completed calibrations, system is operating properly. | No Injury | N/A | During Use in Patient | Yes | 53 | 103
206699 | Unknown fault code displayed | Unknown Fault | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OH | UNITED STATES (USA) | 100267 | Josh the Area manager said the account said they had an issue with their PLS laser and could not use it. No Fault code was given. They were able to complete the case with their other Laser. Patient is doing fine. Please call Josh at 859-338-1186 to let him know when you have scheduled service, he said he can give you a contact and number if you need it. | FSE was not able to duplicate the reported issue, system calibrated and operated properly on arrival. | No Injury | 20231667 | During Preparation for Use | Yes | 45
206759 | The laser system has broken components | Emergency Switch | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100045 | Savanna Campbell the Philips Territory Manager called to report that the OR Department has broken the red e-stop off the back of the PLS Laser. The laser is still working but the plastic red button is broken off the laser.The red estop was broken while the laser was in the equipment storage room.No patient involvement. No case involvement. | FSE verified the red emergency stop button was damaged, replaced. Completed calibrations of catheters and verified proper operation. | No Patient Present | 20231669 | Not during a case | Yes | 45
206768 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100266 | System kept giving fault 106 and 108 during the case.They were able to complete the case with the laser. The patient was doing fine after the case. | FSE found optical rail mis aligned, aligned and completed calibrations. System is operating properly. | No Injury | 20231670 | During Use in Patient | Yes | 45 | 106
206843 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100148 | While doing an Inservice Training the system faulted 106 and error code 250. He noticed that the rear footswitch plug was not working and the red e-stop came off in his hand during the training. He was able to put it back on but he is not sure it is going to stay on.He is requesting an FSE inspect the system. | FSE verified E106 and adjusted the attenuator offset, completed calibrations and system is operating properly. | No Patient Present | 20231675 | Not during a case | Yes | 44 | 106
206860 | Unknown fault code displayed | Unknown Fault | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100215 | Philips Rep Shannon Dockery reported that the PLS Laser faulted while using a Turbo Catheter. He did not know the fault number. He reported that the system seems to fault while using the Turbo Catheters. The system was reset and they completed the case. The patient was doing fine after the case. | FSE determined customer getting errors when lowering fluence below calibration fluence. Rep instructed not to lower below calibration settings, system is operational. | No Injury | 20231677 | During Use in Patient | Yes | 50
206872 | The reference catheter fails to calibrate | Bad Reference Catheter | Unknown | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100090 | The Philips Rep Cary Fennell found that the reference catheter is bad and fails calibration but the case catheters calibrate fine. He is requesting that we send them a new site reference catheter for the system. Send the site a new 2.5 Reference Catheter Part # 425-002.Send to: Carol Ortiz Cath Lab Manager 915-637-0208.Std Fed Ex Overnight delivery. | Non dispatch, replacement catheter was sent to the customer, FSE completed follow up and receipt and operation. | No Patient Present | 20231678 | Not during a case | Yes | 47
206934 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes |  | SPAIN (ESP) | 100004 | Dear Olaf, Diana,today we have had a problem with the PLS in H. Valle Hebron.Dra. Bellera told me that they did the calibration with the non sterile catheter whitout problem, and when they went to calibrate a 0,9mmm Elca catheter the laser machine didn't allow them to do the calibration.They were in a very difficult procedure and they really need the action of the laser, then they decided to try with a 1,4mm Elca catheter, and in that case they were able to calibrate it well and it worked correctly.Dra. Bellera say this is not the first time they have a problem with the calibration of a 0,9mm Elca catheter but they didn't report the issue before because they thought there was a problem with that particular catheter.I'm so sorry but they don't remember what fault number the laser machine showed they only noted that was a calibration issue.During this morning I will be in the hospital. If you want I can try to acces the laser and do some test.I wait for your indications.Thank you so much for your support.Hve a nice day.SandraNOTE: they thrown out the laser catheter box during the procedure and we don't know the serial number of the 0,9mm elca but this morning I will try tto figure out it. | FSE found the optical rail was misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20230313-BV | During Preparation for Use | Yes | 46
206935 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | 06129 | ITALY (ITA) | 100008 | On reviewing logs during PM of today, there is one instance of FC 230 on 29th March 2023 that was not reported. Opening complaint in order to document this occurrence. | During the routine service the FSE noted an E230 in the system log, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230053-BV | Routine Service/Not during a case | Yes | 49 | 230
207215 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100103 | Site Reported that during a case the system gave a fault 106 & 108.They stopped doing the case with the PLS Laser and used CSI to complete the case. The patient was doing fine after. | FSE was not able to duplicate the E106, completed the isolator update, completed calibrations and system is operating properly. | No Injury | 20231106 | During Use in Patient | Yes | 44 | 106
207405 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | FRANCE (FRA) | 100138 | error 310 during a procedure on a patient. A restart of the laser made it possible to complete the procedure without error but this compromised the intervention since after restarting it was necessary to remove the catheter from the patient to recalibrate the catheter. | FSE did not duplicate the reported issue, replaced the footswitch and receptacle as a precaution. System is operating properly. | No Injury | 20230332-BV | During Use in Patient | Yes | 40 | 310
207412 | Clinical catheter fails to calibrate | HVPS | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100287 | FSE received a call from Jovan (philips rep). Stated that the laser would not calibrate during an in service. FSE was close by and went to verify the issue. FSE verified the issue. FSE trouble shot with tech support and determined that it was a footswitch issue.A new foot switch will be sent. | Replacement footswitch was sent to the customer but did not resolve the issue. FSE arrived on site and verified the issue, defective high voltage power supply, system will be deinstalled and replaced with SN 100312. | No Patient Present | 20231712 | Not during a case | Yes | 40
207437 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100104 | The Biomed Billy Balanger called and reported that the Cath Lab told him the laser was not calibrating the catheter in the first few times they tried calibrating. It did finally calibrate in properly. The lab did not tell him what the fault code was it was.They completed the case without any issues, They would like the laser to get checked out before there next case. The patient was doing fine after the case. | FSE found optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20231716 | During Preparation for Use | Yes | 43
207587 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100103 | While doing a PM the customer reported they had a 106 error message on 06/06/23. The patient was treated by other means. No adverse events. Please add a complaint to the call. | During a service the customer reported a E106, not duplicated, the FSE completed the isolator update and calibrations, system is operating properly. | No Injury | 20231106 | During Use in Patient | Yes | 50 | 106
207656 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100287 | Philips Rep Vince Castillo called to report that the new PLS Laser the site received has not worked since they received/installed it. The system footswitch was not working so they were sent new one. The system is giving a system failure 310 and still not working with the new footswitch. He is requesting immediate service on the new PLS Laser. The site is also requesting that there old CVX300 Laser system be returned as a back up for this system. | System has had multiple failures and was deinstalled and replaced with SN 100312. | No Patient Present | 20231712 | During Installation | Yes | 46 | 310
207834 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100289 | Brian Peterson called and reported that the new PLS Laser was giving both Fault 106 and 108 errors during the case. They stopped using the laser and used a Bard Atherectomy Device to complete the case.The patient was doing fine after.FSE Jeff Tobin has already been called by the Philips Rep and is going to the site today. | FSE verified the error and adjusted the attenuator offset and alignments. Completed calibrations and system is operating properly. | No Injury | 20231746 | During Use in Patient | Yes | 42 | 106
207917 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100197 | The PLS laser gave a fault 106 during a case. They stopped using the PLS Laser and continued doing the case using the CVX300 Laser.The patient was doing well after.Call Lab Manager Misty Gerard (239-628-8704) to schedule the service. | FSE verified the error and adjusted the attenuator offset, completed calibrations and system is operating properly. | No Injury | 20231750 | During Use in Patient | Yes | 40 | 106
207928 | The laser system does not provide reliable performance | Debris On Focus Lens | N/A | Alarms/ Messages | Warning/ Message/ Error Message | 30 Day Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100194 | The Philips Rep Jonathan Woods tried to calibrate the PLS Laser. The system fault was No Energy at the Detector. The system was not making any lasing sound. This is a Outpatient Clinic. The Dr decided just to reschedule the patient until the laser was fixed. The patient was open on the table but they had not started anesthesia. There was no additional risk to the patient. The patient was just rescheduled not sent somewhere else.The Philips Rep said that Cedric Benkowski the FSE had already been notified and was on the way to the site.Sheath was in the groin and angiography was performed. Case was peripheral. | FSE found debris on the focus lens and PES sensor, cleaned and replaced sensor. Completed calibrations and system is operating properly. | No Injury | 20231885 | During Preparation for Use | Yes | 85
208269 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100047 | Ryan Schubert the Philips Rep reported that the laser gave a fault 106 during calibration of the laser fiber. The site canceled the case. The patient stayed at the site and was not relocated. The Dr decided to treat the patient medically at this point. The patient was not put at additional risk.Additional information received from rep 20JUN2023:Was the patient open (access)?Was sedation given?Was angiogram performed?RESPONSE:It was groin access, yes to sedation, yes to angiograms performed.Coronary case | FSE was not able to duplicate the reported problem, completed calibrations and system is operating properly. | No Injury,Required Intervention | 20231861 | During Use in Patient | Yes | 34 | 106
208285 | The laser system is damaged by the customer | Environmental Conditions | N/A | Contamination/ Decontamination/ Cleaning | Foreign material | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100079 | David Maydon Philips FSE received a call from Saint David's Hospital for a possible condensation issue at the site. The hospital lost the HVAC system for the entire floor where the laser was kept. They would like the laser checked out for any condensation damage before they use it. | Facility reported loosing HVAC, system was checked without issue and is working properly. | No Patient Present | 20231808 | Not during a case | Yes | 34
208292 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100219 | Jeff Tobin turned in this e-call today from his Philips Territory Rep Hayden Emmington. Hayden emailed Jeff Tobin that a fault 106 had occurred last Friday while doing a case at UC Davis.Hayden called and reported that they were able recalibrate and finish the case using the laser.The patient was doing fine after. They want the laser checked out. | FSE found a noisy PES sensor, completed calibrations and system is operating properly. | No Injury | 20231897 | During Use in Patient | Yes | 35 | 106
208380 | The laser system turns off during a case | Unable to replicate complaint | N/A | Power | Will not turn on | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100154 | Michael Chen Philips Rep called to report that the PLS Laser 100154 had Lost power during a case.They have two PLS Lasers so they just continued the case using the other PLS Laser. The patient was doing fine after the case. | FSE was not able to duplicate the reported problem, system powered up, calibrated, and operated properly. | No Injury | 20231914 | During Use in Patient | Yes | 76
208506 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Catheter Coupler Assy | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100142 | Ace Barte Reported that the cath lab tried to calibrate the 2.3 turbo catheter and it faulted 201. They calibrated the site 2.5 reference catheter fine. They tried to calibrate in the 2.3 catheter again but it faulted 201 again. Ace Barto then gave them a 2.0 Turbo catheter which he had with him and it calibrated in fine and they were able to do the case without further issues. The patient was doing fine after.Fault 201 is Invalid Catheter ID Error. | FSE found a defective coupler switch PCB, replaced. completed calibrations and system is operating properly. | No Injury | 20231920 | During Preparation for Use | Yes | 40 | 201
208508 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100228 | Bri Berardi Philips Rep reported that the system gave a fault 106 during the case. They recalibrated and were able to complete the case without further issue. The patient was doing fine. | FSE was not able to duplicate, completed calibrations and alignments, system is operating properly. | No Injury | 20231927 | During Use in Patient | Yes | 47 | 106
208706 | The laser system will not turn on | Interlocks Intermittently Dis-Engage | N/A | Power | Will not turn on | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100223 | Dierdre Tirino Philips Rep was trying to do an in-service on the laser but it would not turn on. She tried flipping the main breaker and resetting the red kill switch to no avail. The system will not power up. | FSE found rear panel loose, tightened and engaged interlocks properly. System is operating properly. | No Patient Present | 20231931 | Not during a case | Yes | 46
208718 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 303 | Fault 303 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Fl | UNITED STATES (USA) | 100179 | Tony Daley FSE reported that Terra from the site contacted him to report that the laser gave a fault 303 and 108 last week. The case was completed with the laser. There were no other issues. The patient was fine after. | FSE was not able to duplicate, completed calibrations and system is operating properly. | No Injury | 20231596 | During Use in Patient | Yes | 46 | 303
208738 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100219 | Our Territory Manager Hayden called to report that while using 0.9mm ELCA that the system faulted Error 106 after increasing settings to 80/60.Hayden said the case was completed with laser and patient is doing well | FSE verified E106, adjusted attenuator offset and completed calibrations. System is operating properly. | No Injury | 20231947 | During Use in Patient | Yes | 43 | 106
208809 |  |  | Fault 600 |  |  |  | Closed | No |  | ITALY (ITA) | 100241 | During a training session on the use of PLS, the laser was turned on. After the warmup, the calibration was initiated with a demo sheath. In the middle of the calibration an error message appeared “error 600, alert on temperature” We tried to turn off and turn on again the laser but every time the message “error 600” appeared. After 30 minuts, we tried to turn on again the laser and, this time, it was possible to complete the calibration but after that there was again the error message 600. |  | No Patient Present | 20230371-BV | Not during a case |  | 1 | 600
208810 |  |  | Fault 600 |  |  |  | Closed | No |  | ITALY (ITA) | 100241 | During a training session on the use of PLS, the laser was turned on. After the warmup, the calibration was initiated with a demo sheath. In the middle of the calibration an error message appeared “error 600, alert on temperature” We tried to turn off and turn on again the laser but every time the message “error 600” appeared. After 30 minuts, we tried to turn on again the laser and, this time, it was possible to complete the calibration but after that there was again the error message 600. |  | No Patient Present | 20230371-BV | Not during a case |  | 1 | 600
208811 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | ITALY (ITA) | 100241 | During a training session on the use of PLS, the laser was turned on. After the warmup, the calibration was initiated with a demo sheath. In the middle of the calibration an error message appeared “error 600, alert on temperature” We tried to turn off and turn on again the laser but every time the message “error 600” appeared. After 30 minuts, we tried to turn on again the laser and, this time, it was possible to complete the calibration but after that there was again the error message 600. | FSE verified E600 and disconnected, reconnected fuses to the DC isolator, completed calibrations and system is operating properly. | No Patient Present | 20230371-BV | Not during a case | Yes | 51 | 600
208880 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes |  | SPAIN (ESP) | 100004 | Dear all,Dra. Bellera from H. Valle Hebron (interventional cardiology department) just call me because they can not do the calibration with the 0.9mm Elca catheter (they can calibrate with the test catheter but not with the sterile one).I don't have more details because the patient is still on the table.This time the doctor is aware she has to keep the elca catheter.This is a very delicate situation, is the second time that happen in this hospital.We need to do a service as quick as we can.Abdelkarim,could you send me the service report you did the last time, please?The biomed department is asking for it.I will tell you more details when I can talk with the doctor again.Kregards.SandraADDITIONAL INFORMATION OBTAINED 04JUL2023:The doctor could not start the treatment with the laser because the 0’9mm sterile catheter didn’t work.Due this the patient was treated with other technique and the procedure finished fine. The patient was discharged several days after. | FSE found optical rail misaligned, replaced the focus lens and aligned. completed calibrations and system is operating properly. | No Injury | 20230375-BV | During Use in Patient | Yes | 41
208916 |  |  | Fault 600 |  |  |  | Closed | No | OR | UNITED STATES (USA) | 100198 | While doing an Inservice training the laser started giving a fault 600 High temp alert. They turned the system off and let it cool overnight. The next morning the laser was still giving a fault 600 when they turned it back on. |  | No Patient Present | N/A | Not during a case |  | 0 | 600
208917 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OR | UNITED STATES (USA) | 100198 | While doing an Inservice training the laser started giving a fault 600 High temp alert. They turned the system off and let it cool overnight. The next morning the laser was still giving a fault 600 when they turned it back on. | FSE found DC cables routing caused an intermittent disconnect, adjusted routing and system is operating properly. | No Patient Present | 20231973 | Not during a case | Yes | 41 | 600
208938 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CO | UNITED STATES (USA) | 100300 | Jenn Dowell Philips Rep in the case reported that the laser faulted 106 twice. They then tried a 1.4 Turbe catheter which also failed and faulted 106. | FSE verified and adjusted the attenuator offset and alignment. Completed calibrations and system is operating properly. | No Injury | 20231972 | During Use in Patient | Yes | 44 | 106
209179 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | SC | UNITED STATES (USA) | 100165 | Error Codes 106 found during the review of the System Log on 30JUN2023 | During service the customer reported a 106, completed alignments and calibrations, system is operating properly. | No Patient Present | 20231184 | Routine Service/Not during a case | Yes | 42 | 106
209543 | Clinical catheter fails to calibrate | Debris On Focus Lens | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | CT | UNITED STATES (USA) | 100074 | Laser failed to calibrate both the 2.5mm catheter and 1.7mm clinical catheter. Rep attempted several times to calibrate without success. Case was not completed with the laser. It was completed with a Medtronic atherectomy device without delay at the same facility. | FSE found debris on the focus lens, cleaned. Completed calibrations and system is operating properly. | No Injury | 20232020 | During Preparation for Use | Yes | 61
209592 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100217 | E600 during a case. Rep restarted the laser multiple times and completed the case with the laser at the same facility without delay. | FSE was not able to duplicate the reported E600, completed calibrations and system is operating properly. | No Injury | 20232024 | During Preparation for Use | Yes | 56 | 600
209597 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IN | UNITED STATES (USA) | 100239 | E600 occurred at the end of the case. Case completed with the laser without delay at the same facility. | FSE was not able to duplicate the reported problem, reseated fuse connectors in the DC isolator circuit. Completed calibrations and system is operating properly. | No Injury | 20232026 | During Use in Patient | Yes | 60 | 600
209742 | The laser system turns off during a case | System Turned Off | N/A | Power | Other | Non-Reportable | Closed | Yes |  | FRANCE (FRA) | 100292 | PLS start and then shut down. Restart IMPOSSIBLE. | FSE found a poor connection in the keyswitch/e stop cable, repaired. Completed calibrations and system is operating properly. | No Injury | 20230390-BV | During Preparation for Use | Yes | 52
209934 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100181 | error 500 | FSE found a defective shutter, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20230395-BV | Not during a case | Yes | 56 | 500
209935 |  |  | Fault 500 |  |  |  | Closed | No |  | ITALY (ITA) | 100181 | error 500 |  | No Patient Present | 20230395-BV | Not during a case |  | 1 | 500
209977 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MI | UNITED STATES (USA) | 100049 | Dennis Keast Philips Rep emailed Tech Support a pic of the PLS Laser Temp Error 600. Dennis Keast has not returned my teams attempt to talk to him or returned any messages. I called and talked Sherie Scholtz EP Lab Manager about the Error 600. Sherie reported that they were doing an in-service at the time of the error 600. The EP Dept was being trained on the PLS system. No Patient was on the table. She also reported that this is the Cath Labs Laser. They use it occasionally for lead extraction cases. | FSE found a defective PCN, customer is billable and working on entitlement. | [5 empty] | 600
209982 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100264 | Bayron Murillo FSE Philips received a call from Onsite Lab Manager Lisa McGovern that the laser gave a fault 106 during a case. They were using the .9 catheter at 50/50 when the fault occurred. They then completed the case with a balloon. The patient was fine after the case. | FSE adjusted the attenuator offset one step, completed calibrations and system is operating properly. | No Injury | 2022044 | During Use in Patient | Yes | 56 | 106
210093 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NV | UNITED STATES (USA) | 100290 | Dave Patel CCL Mgr reported to Philips Rep Darren Dennings that the new PLS Laser gave a fault 106 during a case with the .9 Catheter.Dave reported that it started working after restating and recalibrating and using a 2nd .9 catheter. They decided to finish this case using the Roto blader device. The patient was fine after the case. | FSE verified the E106, found optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20232604 | During Use in Patient | Yes | 73 | 106
210102 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100298 | Gerardo Valencia Philips Rep reported that during a case the PLS Laser faulted 106 with the .9 catheter at 80/80. They restarted and recalibrated the .9 fine but when they went to 80/80 again it faulted 106. They stopped using the laser and then completed the case with a balloon. The patient was fine after.Contact in the OR Department is Ashley Washington 817-675-6936. | FSE verified E106, completed calibrations and alignments, system is operating properly. | No Injury | 20232056 | During Use in Patient | Yes | 55 | 106
210225 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100293 | E600 when turning on laser. Case completed with angioplasty | FSE verified the E600, disconnected and reconnected the DC isolator fuses, completed calibrations and system is operating properly. | No Injury | 20232214 | During Preparation for Use | Yes | 58 | 600
210263 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OH | UNITED STATES (USA) | 100314 | When going above settings 60/60 laser gives a E106. Case was completed with laser at lower settings | FSE was not able to duplicate the reported problem, noted vessel alignment was out slightly, aligned. Completed calibrations and system is operating properly. | No Injury | 20232216 | During Use in Patient | Yes | 58 | 106
210267 | Unknown fault code displayed | Unable to replicate complaint | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100038 | Unknown fault. Laser was reset and fault cleared. Case completed with laser | FSE was not able to duplicate the reported problem, did note an E108 and E221 in the log. Completed calibrations and system is operating properly. | No Injury | 20232218 | During Preparation for Use | Yes | 58
210339 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100160 | Customer (Kris Ramboying) reported to Kim Adams Philips Rep that the laser faulted 106 on calibration. They were able to recalibrate and complete the case without problems. This happened on two separate cases the same fault code 106.The patients of both cases were fine after. | FSE verified E106, installed isolators. Completed calibrations and system is operating properly. | No Injury | 20232223 | During Preparation for Use | Yes | 50 | 106
210687 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100324 | Lab Manager Claribel Ramirez called Philips Rep Matt Tulgan to report that the PLS Laser gave a fault 106 during a case with the .9 catheter. They stopped using the laser and finished the case with a balloon. The patient was fine on completion of the case. | FSE found a noisy PES, replaced and also replaced the HMI due to a screen failure. Completed calibration and system is operating properly. | No Injury | 2021687 | During Use in Patient | Yes | 49 | 106
210694 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100219 | Jeff Tobin FSE received an email from Philips Territory Mgr Darren Dennings that the laser gave a fault code 106 while using the .9 catheter at a high rep rate.We do not know how the case was completed. (Additional information received on 31JUL2023: case was completed with roto and stent).They did report that the patient was fine on completion of the case. | FSE found the optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20232277 | During Use in Patient | Yes | 60 | 106
210832 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100160 | Customer (Kris Ramboying) reported to Kim Adams Philips Rep that the laser faulted 106 on calibration. They were able to recalibrate and complete the case without problems. This is the second time this has happened during a case. The patient was fine on completion of the case. | FSE completed the isolator update on the system PES and VES, completed calibrations and system is operating properly. | No Injury | 20232223 | During Preparation for Use | Yes | 59 | 106
210912 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 501 | Fault 501 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100228 | Bri said they got an Error code 501, reset laser power several times but could not get past the Error code. Case was completed with CSI and patient is doing well. Bri Cucci (TM) cell # 561-212-1977 | FSE found a defective shutter, replaced. Completed calibrations and system is operating properly. | No Injury | 20232300 | During Preparation for Use | Yes | 58 | 501
210940 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | INDIA (IND) | 100076 | Fault Code:106 with 0.9mm catheter while testing system to prepare system for a schedule case on 28th july2023. Today customer was testing system as a preliminary check without patient and noticed error 106. | FSE found the optical rail misaligned and debris on the optics, aligned and cleaned. Completed calibrations and system is operating properly. | No Patient Present | 20232304 | Not during a case | Yes | 61 | 106
211128 | The power cord cable is damaged | Damaged Power Cord | N/A | Hardware | Mechanical issue (except movement, structure, or table) | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100159 | Wade said they moved the laser while it was plugged in and damaged the power cord. No case was involved. Please ship overnight a new power cord for a PLS laser to Amanda Stanley 405-760-2283They will need this for cases on Monday morning. Wade also asked that we not bill the customer for shipping on this order. He said that has been happening lately. | Replacement power cord was sent to the customer, FSE completed follow up on receipt and proper operation. | No Patient Present | N/A | Not during a case | Yes | 56
211166 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IL | UNITED STATES (USA) | 100258 | Jeff said the laser gave a 103 error code. They could not get past it so they completed the case with a competitor's Laser device. Patient is doing well. Rep Jeff Sutter @ 719-641-3504 | FSE was not able to duplicate the reported problem, completed calibrations and system is operating properly. | No Injury | 20232326 | During Use in Patient | Yes | 59 | 103
211549 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Hardware | Display / Monitor issue | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100324 | Upon power on, the HMI has a blue screen with white lines going through it. Customer reset the laser multiple times and the HMI cleared. Case was completed with the laser without delay at the same facility. | FSE found a defective HMI, replaced. Completed calibrations and system is operating properly. | No Injury | 20232369 | During Preparation for Use | Yes | 50
211564 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100324 | E106 during the case. Case was not completed with the laser. Case was completed with a balloon with no delay at the same facility. | FSE verified error and completed calibrations and alignments, system is operating properly. | No Injury | 20232276 | During Preparation for Use | Yes | 50 | 106
211636 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100098 | During an in-service training, E106 occurred multiple times with 3 different catheters. No patient present. | FSE verified and adjusted the attenuator offset, completed calibrations and system is operating properly. | No Patient Present | 20232373 | Not during a case | Yes | 49 | 106
211660 |  |  | N/A |  |  |  | Closed | No | CT | UNITED STATES (USA) | 100074 | Sales Rep states that there was an odd sound coming from the laser as well as a burning smell. Case was completed with the laser without delay at the same facility. |  | No Injury | N/A | During Use in Patient |  | 0
211661 | The laser system exhausts excessive heat | Low Max Energy < 3 Months | N/A | Hardware | smoke, fire, or smokey odor | Non-Reportable | Closed | Yes | CT | UNITED STATES (USA) | 100074 | Customer states that there was an odd sound coming from the laser as well as a burning smell. Case was completed with the laser without delay at the same facility. | FSE found low energy, system was deinstalled and replaced with SN 100346. | No Injury | 20232376 | During Use in Patient | Yes | 49
211879 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100335 | Upon calibration of the .9 Elca Catheter the laser displayed Failed Calibration/ Recalibrate. They opened two more .9 catheters and all three catheters failed calibration. Andy Watts the Philips Rep in the case was able to calibrate in the site reference catheter without a problem. They were able to go and get the CVX300 Laser and use the .9 catheters that had just failed calibration on the PLS Laser.The patient was fine on completion of the case. | FSE found optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20232390 | During Preparation for Use | Yes | 45
211974 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | WA | UNITED STATES (USA) | 100272 | Site Tech Stephanie Thurren called to report that the PLS Laser was giving an Error Code of 230 while calibrating the .9 catheter. They tried restarting and recalibrating the .9 catheter, but the system gave an error code of 230 again. The patient was given a IVUS diagnostic. The Dr. decided that no other treatment needed to be done. The patient may be brought back at another time. Patient was fine after completion of the IVUS diagnostic. | FSE found a defective HVPS, replaced. Completed calibrations and system is operating properly. | No Injury | 2032394 | During Preparation for Use | Yes | 44 | 230
211976 | The laser system has broken components | Damaged Footswitch | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100333 | Paul Guidry Clinical Philips Rep called to report that the Key assembly on the laser had been broken and laying on the ground. The laser is unusable. They still have a CVX300 Laser as a backup to use for cases until the PLS Laser has been fixed. Paul has already called the FSE Don Balzi who is in route. | FSE verified damaged keyswitch, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20232395 | Not during a case | Yes | 44
212116 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100282 | Site had calibrated in the case catheter fine and was almost done when the system faulted 103. The Dr was very happy with the results of the laser. No other device was needed. The patient was fine on completion of the case. They notified the Philips Rep William Follensbee that Philips needed to send an FSE to look at the new PLS Laser. | FSE verified the error and completed calibrations and alignments. System is operating properly. | No Injury | 20232405 | During Use in Patient | Yes | 42 | 103
212132 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | SC | UNITED STATES (USA) | 100192 | Please process a complaint for Coastal A#94399828. Laser catheter 414-151 lot FAY23D21A. It did not calibrate during prep and outside the body. They opened a second one and it was successful. No harm to patient. CF Access. Therapeutic case. Case date was 8/10/23. VS Dr. Sean Hislop. I was notified today and was not there for the procedure. 106 error code. | FSE was not able to duplicate the reported 106, completed alignment and calibrations and the system is operating properly. | No Injury | 20232410 | During Use in Patient | Yes | 42
212133 |  |  | N/A |  |  |  | Closed | No | SC | UNITED STATES (USA) | 100192 | Please process a second complaint (different case) for Coastal A#94399828. Laser catheter 414-151 lot FAY23D21A. It did not calibrate during prep and outside the body. They opened a second one and it was successful. No harm to patient. CF Access. Therapeutic case. Case date was 8/10/23. VS Dr. Sean Hislop. I was notified today and was not there for the procedure. Discarded and not returning. |  | No Injury | N/A | During Use in Patient |  | 5
212197 | Clinical catheter fails to calibrate | Bad Reference Catheter | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100141 | Can not calibrate the laser. Multiple test catheters were used. Second laser on site was used during the case | FSE found a defective reference catheter, replaced. Completed calibrations and system is operating properly. | No Injury | 20232412 | During Preparation for Use | Yes | 39
212210 |  |  | Fault 106 |  |  |  | Closed | No | TN | UNITED STATES (USA) | 100282 | I have complaint call for e230 on sn 100272, 20232394, in-progress. I completed the work last Thursday. For all my hard work, I was rewarded with sales rep Todd Christensen reporting that they had multiple e106 on Friday. They successfully calibrated a 0.9mm and observed a e106 at 80/80. They re-calibrated and observed the same result. They tried a second 0.9mm, and observed e106 at 60/80. At this point the physician decided to use SHOCKWAVE! In liu of the laser. Patient was fine after. |  | No Injury | N/A | During Use in Patient |  | 1 | 106
212211 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | WI | UNITED STATES (USA) | 100272 | I have complaint call for e230 on sn 100272, 20232394, in-progress. I completed the work last Thursday. For all my hard work, I was rewarded with sales rep Todd Christensen reporting that they had multiple e106 on Friday. They successfully calibrated a 0.9mm and observed a e106 at 80/80. They re-calibrated and observed the same result. They tried a second 0.9mm, and observed e106 at 60/80. At this point the physician decided to use SHOCKWAVE! In liu of the laser. | FSE verified the E106, during troubleshooting the system failed with an E230. FSE found the system has a PCN failure and will be deinstalled and replaced with SN100342. | No Injury | 20232419 | During Use in Patient | Yes | 50 | 106
212301 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100094 | Justin Pisciotta Philips Rep called Ken Keao the Philips FSE and reported that the site had a fault 106 happen during the case. They recalibrated the catheter and were able to complete the case with the laser. The patient fine on completion of the case. | FSE was not able to duplicate the reported E106, completed calibrations and alignments and system is operating properly. | No Injury | 20232472 | During Use in Patient | Yes | 52 | 106
212455 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 104 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100200 | The Philips Rep Rachell Bawroski called to report that she had been notified by the site that the PLS laser faulted 104 or 106 during calibratrion. They shut off the failed PLS Laser and used the backup CVX300 to complete the case. They were able to use the same catheter with the CVX300 that they had opened for the failed PLS Laser.The patient was fine on completion of the case. | FSE verified the error in the log, completed calibrations and alignment, system is operating properly. | No Injury | 20232573 | During Preparation for Use | Yes | 50 | 104
212609 | The laser system will not turn on | Wiring | N/A | Power | Will not turn on | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100158 | Sales Rep Ryan Delucas @ 856-217-2491 said he could not get the laser to turn on, Checked Estop, Circuit breaker and Power cord. Case was completed by the Pheonix atherectomy device and patient is doing well. | FSE found a loose connection on the PIM, tightened all connection. Completed calibrations and system is operating properly. | No Injury | 20232603 | During Preparation for Use | Yes | 46
212615 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NV | UNITED STATES (USA) | 100290 | System had Error 106 occur. After many tries using a few different catheterrs we were able to complete case. | FSE found a noisy PES sensor, replaced. completed calibrations and system is operating proprely. | No Injury | 20232604 | During Preparation for Use | Yes | 65 | 106
212617 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Catheter Coupler Assy | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | WA | UNITED STATES (USA) | 100342 | When attempting to install SN 100342 I observed E201. It required corrective action. | FSE found catheter coupler A34 PCB damaged on installation, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20232424 | During Installation | Yes | 46 | 201
212650 | The laser system indicates fault codes 300-399 - State faults (NEX) | Electro / Mechanical | Fault 301 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100033 | Paul Costello (TM) said the laser gave a 301 error code. Patient is well waiting to hear back from Paul as to how the patient was treated. Update: Paul said they did use the laser to complete the case, They used a 2.0 laser fiber to finish treating the patient.Ref customer complaint 082123 | FSE confirmed error was actually a 103, completed calibrations and alignments and system is operating properly. | No Injury | 20232606 | During Use in Patient | Yes | 45 | 301
212783 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100248 | Mike said they got an error code 103 during the case. Case was completed with Balloon and patient is doing well. | FSE verified and adjusted optical rail alignment and calibrations. System is operating properly. | No Injury | 20232614 | During Use in Patient | Yes | 51 | 103
212858 | Multiple Fault Codes Displayed | Shutter Assembly | Fault 210 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100228 | Bri said the laser has error code 210 and 500 . rebooted several times but could not get past the error codes. Case was completed with CSI device. Patient is doing well. | FSE found a defective shutter, replaced. Completed calibrations and system is operating properly. | No Injury | 20232621 | During Preparation for Use | Yes | 53 | 210
212863 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IN | UNITED STATES (USA) | 100239 | Sofia said they got the 600 error code during the case , let the laser cool down for a few minutes and lowered the catheter setting to 50/50 and was able to complete the case with the laser. Patient is doing well. | FSE found the sheathing was preventing the fuses from a proper connection, repaired. System is operational. | No Injury | 20232624 | During Use in Patient | Yes | 53 | 600
212865 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IN | UNITED STATES (USA) | 100235 | Sophia said while she was working a case in one room their other PLS laser was being used in another room and they also got an error code 600. They powered down the laser and rebooted and recalibrated the catheter and was able to complete the case. Patient is doing well. | FSE found the cable sheathing was preventing the fuse holders to make contact, repaired. System is operational. | No Injury | 20232623 | During Use in Patient | Yes | 53 | 600
212872 | Unknown fault code displayed | Optical Rail Mis-Aligned | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100254 | Mike said they got an error during the case, retried with a second catheter as was able to complete the case. They did not get the error code when this happened. patient is doing well. | FSE found the optical rail mis aligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20232626 | During Use in Patient | Yes | 53
212879 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100200 | Error Code 221 found during the review of the System Log on 3/30/23 @ 19:16 | During service the FSE noted a 221 in the system logs, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20232573 | Not during a case | Yes | 50 | 221
212880 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100200 | Error Code 106 X3 found during the review of the System Log on 6/30/23 @ 11:35 | During service the FSE noted a E106 in the system logs, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20232573 | Not during a case | Yes | 50 | 106
212881 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 252 | Fault 252 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100200 | Error Code 252 found during the review of the System Log on 7/27/23 @ 11:36 | During service the FSE noted a E252 in the system logs, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20232573 | Not during a case | Yes | 50 | 252
212965 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Missouri | UNITED STATES (USA) | 100137 | From Amy - Catheter was prepped and calibrated correctly. With message of "Cal Okay"Fluency and Rate were set at 45/45 and the "set to Ready" button was selected once catheter had been advanced over the wire into the body. the pedal is activated and within 30 seconds the system turns off with an error message of "106.' The Nexcimer console was turned off and back on and the catheter was re-calibrated. The physician did not like that there was no red light was seen at tip of catheter so he asked for a second catheter to be opened and the first catheter to be replaced. case was completed with laser and second catheter. Patient is doing well. this is linked to a catheter complaint 211263 and we notified by PMS about this today (8/25/2023) to open the laser side of this complaint. The original catheter complaint came in on 7/26/2023 which is also when the 106 error happened. | FSE was not able to duplicate the E106, completed alignments and calibrations and system is operating poperly. | No Injury | 20232632 | During Preparation for Use | Yes | 49 | 106
213066 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MN | UNITED STATES (USA) | 100316 | FSE reported - Rep Sam Akins called to report that on Friday they observed an e106 with a 2.0mm te during a PV case. The error occurred on the last pass with the laser, and the physician was satisfied with the laser results. | FSE was not able to duplicate the reported problem. Completed alignment and calibrations and system is operating properly. | No Injury | 20232642 | During Use in Patient | Yes | 46 | 106
213254 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | Fault 102 | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100181 | Dear Diana,can you please open a Work Order for the PLS 100181?Our distributor called me reporting the following issue:Calibration fail during 0.9 mm catheter calibration.I had already a chat with Al, here in copy. We need to the visit customer asap.I’ll be in Catania on Monday 4th in the morning.Prior to the FSE's assistance, two ELCA were rightly pulled away from the packaging, prepared and rightly connected to the laser. After the connection an error appear to the laser, so they change the catheter but the problem was the same. Physician was unable to perform measurment with the laser.04 of september our FSE performed the assistance.ERROR 102Thanks,Kind Regards,Martino | FSE found optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20230518-BV | During Preparation for Use | Yes | 44 | 102
213255 |  |  | N/A |  |  |  | Closed | No |  | ITALY (ITA) | 100181 | Dear Diana,can you please open a Work Order for the PLS 100181?Our distributor called me reporting the following issue:Calibration fail during 0.9 mm catheter calibration.There’s no error codeNo patient involvedI had already a chat with Al, here in copy. We need to the visit customer asap.I’ll be in Catania on Monday 4th in the morning.Thanks,Kind Regards,Martino |  | No Patient Present | 20230518-BV | Not during a case |  | 6
213274 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | NC | UNITED STATES (USA) | 100215 | Coupler PIN code test fails due to bad part. (missing spring contact) | During routine service the FSE noted a defective coupler A34, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20232173 | Routine Service/Not during a case | Yes | 44
213279 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100099 | Customer experienced an E108 and rep was able to talk through getting it cleared and up and running. Customer was able to finish the case without further issue. | Customer experienced a E108, due to depressing footswitch too quickly after ready selection. System is operating properly. | No Injury | N/A | During Preparation for Use | Yes | 44 | 108
213387 | Multiple Fault Codes Displayed | Fault 106 | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AR | UNITED STATES (USA) | 100304 | Rep called to report an E103 on initial pass with a 2.0mm catheter and then after recalibration the system displayed a 106. Case was completed by other means at the same facility and time. | FSE verified E106, adjusted atten offset, alignment and calibrations. System is operating properly. | No Injury | 20232673 | During Use in Patient | Yes | 43 | 103
213616 | The laser system generates unusual noise | Unable to replicate complaint | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100205 | Upon powering on the laser, the interlock buzzer is alarming and cannot be silenced. Case was not completed with the laser. It was completed with their CVX without delay at the same facility. | FSE was not able to duplicate the reported problem, system power up and operated normally on arrival. | No Injury | 20232695 | During Preparation for Use | Yes | 44
213695 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100199 | Catheter coupling missing #1 contact clip. Replaced this part. | During the routine service the FSE noted the catheter coupler A34 damaged, replaced. completed calibrations and system is operating properly. | No Patient Present | 20231957 | Routine Service/Not during a case | Yes | 57
213724 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CT | UNITED STATES (USA) | 100346 | E106 after calibrating the clinical catheter. Case was not completed with the laser. It was completed by other means without delay at the same facility. | FSE found a loose connection and adjusted atten offset, completed calibrations and system is operating properly. | No Injury | 20232702 | During Use in Patient | Yes | 62 | 106
213815 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100081 | E103/Calibration Error with a 0.9mm coronary catheter. Case was not completed with this laser. It was completed with their CVX300 without delay at the same facility. | FSE was not able to duplicate the reported problem, completed calibrations and system is operating properly. | No Injury | 20232712 | During Preparation for Use | Yes | 61 | 103
213824 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100323 | E230 after the 4th pass with the laser. Case was not completed with the laser. It was completed with angioplasty without delay at the sane facility. | FSE verified error, seated connections and cycled power to clear. Completed calibrations and system is operating properly. | No Injury | N/A | During Use in Patient | Yes | 61 | 230
213908 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IN | UNITED STATES (USA) | 100239 | E106 occurred during 0.9 catheter at 30/25 settings. System was restarted and catheter calibrated and run at 40/40 with no issues. Case was continued with laser and completed with laser at the same facility without delay. | FSE verified and adjusted atten offset, completed calibrations and system is operating properly. | No Injury | 20232726 | During Preparation for Use | Yes | 67 | 106
214005 | Clinical catheter fails to calibrate | Unable to replicate complaint | Unknown | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100023 | The laser is not calibrating a 0.9 cath but is calibrating the 1.4 and Reference catheters. Case was completed with 1.4 catheter. Patient is doing well. No Error code seen. | FSE was not able to duplicate the reported problem, system calibrated and operated properly on arrival. Did note replacement of coupler due to cracked A34. | No Injury | 20232731 | During Preparation for Use | Yes | 57
214055 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100003 | Fault 108 during operation with a 14F GlideLight catheter, Lot- FGB23B28ACase completed with TightRail 13F. System restarted correctly when switched off/on | FSE instructed account on E108 which is caused by user and system is operating properly. Non dispatch. | No Injury | 20230544-BV | During Use in Patient | Yes | 56 | 108
214462 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100312 | Sales rep Vincent 813-720-0718 Said they got a Fault code 106 and 105 while trying to calibrate a catheter. Reset power to laser and was able to get past the codes and complete the case. Patient is doing well. | FSE found optical rail misalignment, aligned and completed calibrations, system is operating properly. | No Injury | 20232758 | During Preparation for Use | Yes | 78 | 106
214839 | The laser system has malfunctioning components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100361 | Replacement of a coupler during installation. | During installation the FSE noted a defective A34 fiber coupler PCB, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20232581 | During Installation | Yes | 70
214849 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100201 | Laser gives E106 when trying to use it | FSE verified and adjusted the attenuator offset by one step, completed calibrations and system is operating properly. | No Injury | 20232942 | During Preparation for Use | Yes | 70 | 106
214929 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Unable to replicate complaint | Fault 200 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100061 | Rep Mike Carney reported an E200, removed catheter and reinserted which worked but then the system would not calibrate. Removed and reinserted catheter getting a E200 again, removed and reinserted and system calibrated and completed the case without issue. | FSE was not able to duplicate the reported problems, system operated properly on arrival. | No Injury | 20232955 | During Preparation for Use | Yes | 69 | 200
215133 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100383 | The VES sensor is bad on SN 100383 at Woodland Heights MC in Lufkin during a new install. | FSE noted a defective VES at installation, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20232735 | During Installation | Yes | 66
215244 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100384 | An actuator lever broke in the fiber coupler assembly and interfered with other catheter levers. Problem has been resolved with another fiber coupler from trunk spares. | During installation the FSE noted a damaged fiber coupler, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20232954 | During Installation | Yes | 65
215355 | Multiple Fault Codes Displayed | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100357 | Rep called to report an E600 and E108, case was completed by other means at the same facility and time. | FSE was not able to duplicate reported problem, removed zip ties to relieve stress on fuse holders. Completed calibrations and system is operating properly. | No Injury | 20233019 | During Use in Patient | Yes | 72 | 600
215488 | Clinical catheter fails to calibrate | Operator Error | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100104 | Two different 1.7mm catheters failed to calibrate on multiple tries. When calibration succeeded the DR pressed the footswitch and an unknown error message occurred. After a 10 second countdown they tried again and successfully completed the case with laser. | Determined the problem was due to operator error not releasing the footswitch during wait time. System is operating properly. | No Injury | 20233030 | During Preparation for Use | Yes | 63
215489 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | HI | UNITED STATES (USA) | 100177 | E106 when trying to use the laser | FSE verified and updated software per CR, completed calibrations and system is operating properly. | No Injury | 20233031 | During Preparation for Use | Yes | 63 | 106
215518 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | NETHERLANDS (NLD) | 100018 | PLS laser system. (picture attached to same email along with the complain).Fault code 106 after preparing the catheter on the laser system.Now the system is working well (Reboot the system and after reboot system was ok) but when there is a PM please check if the problem can be solved. During an ELCA case this message pops up when we wanted to start lasering. | FSE verified and updated software per CR, completed calibrations and system is operating properly. | No Patient Present | 20230581-BV | Not during a case | Yes | 62 | 106
215644 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100392 | Fiber coupler board is broken, pin 1 lever broke off. Please order 0101-0348 A34 board quant. 2 for tomorrow morning against this call. | During installation the FSE noted a damaged fiber coupler A34, replaced the coupler and completed calibrations, system is operating properly. | No Patient Present | 20232989 | During Installation | Yes | 59
215648 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OR | UNITED STATES (USA) | 100288 | E106 during the case with a 0.9mm catheter. The sales rep had the account increase the fluence with a new catheter and the case was completed with the laser without delay and at the same facility. | FSE verified and completed SW update per CR, Completed calibrations and system is operating properly. | No Injury | 20233039 | During Use in Patient | Yes | 65 | 106
215719 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100374 | E106 during the case. Case was not completed with the laser. It was completed by other means without delay at the same facility. | FSE found a defective PES with a noisy signal, replaced. Completed calibrations and system is operating properly. | No Injury | 20233041 | During Use in Patient | Yes | 64 | 106
215789 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100234 | 106 error codeNo patient on the table | FSE completed software update per CR, completed calibrations and system is operating properly. | No Patient Present | 20230583-BV | Not during a case | Yes | 65 | 106
215790 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100016 | E500 at startup that would not clear with multiple reboots of the laser. Case was completed with their 2nd laser without delay at the same facility. | FSE found a defective shutter, replaced. Completed calibrations and system is operating properly. | No Injury | 20233048 | During Preparation for Use | Yes | 65 | 500
215800 | The laser system covers are damaged | Damaged Laser System Cover | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100298 | Customer called to report that the front cover is broken. Case was completed with the laser without delay at the same facility. | FSE verified damaged cover and replaced. System is operational. | No Injury | 20233053 | During Preparation for Use | Yes | 65
215909 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100249 | E600 at startup. Case not completed with the laser. Case was completed with a balloon without delay at the same facility. | FSE verified the E600, inspect and reconnect fuse connections to DC isolator, completed calibration and system is operating properly. | No Injury | 20233068 | During Preparation for Use | Yes | 64 | 600
215915 | Clinical catheter fails to calibrate | Debris On Focus Lens | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100247 | Would not calibrate a catheter. Case was not completed with this laser. the case was completed with their 2nd laser without delay at the same facility. | FSE found dust on the rail optics, cleaned and completed calibrations, system is operating properly. | No Injury | 20233074 | During Preparation for Use | Yes | 64
215921 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100117 | E106 during the case with a 0.9mm catheter. Case was not completed with this laser. Case was completed with a balloon without delay at the same facility. | FSE completed the software update per CR, completed calibrations and system is operating properly. | No Injury | 20233077 | During Use in Patient | Yes | 64 | 106
215984 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100141 | E106 with two 0.9mm catheters. Laser was reset, catheters recalibrated and the case was completed with the laser without delay at the same facility. | FSE completed software update per CR, completed calibrations and system is operating properly. | No Injury | 20233078 | During Use in Patient | Yes | 63 | 106
216158 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100254 | Bryan said the account got unable to calibrate error while trying to calibrate several clinical catheters. They rebooted the laser all with no help. Patient was treated with Balloon and Stent and is doing well. | FSE was not able to duplicate the reported problem, adjusted alignment and completed calibrations and system is operating properly. | No Injury | 20233087 | During Preparation for Use | Yes | 69
216244 | The footswitch connection is not functioning | Damaged Footswitch | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | MN | UNITED STATES (USA) | 100037 | Brent said they are having issues with the footswitch being loose and intermitting not working. They were able to unplug it and plug it back in and it worked for the case today. Patient is doing well. | FSE verified a damaged footswitch, replaced. Completed calibrations and system is operating properly. | No Injury | 20233102 | During Preparation for Use | Yes | 71
216595 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Tokyo | JAPAN (JPN) | 100331 | Error code 221 occurred when the system was turned on. The PLS was updated to ver1.0.1.0 during warehouse incoming inspection but no error code 221 occurred then. FSE Homura will go on-site and exchange attenuator motor that BIU provided for system ver update.Was the Device in Clinical Use at the Time the Issue was discovered? -> NoWas any Patient or User Harmed? -> NoBecause the system cannot output laser when calibration is not completed. | FSE found a defective attenuator, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20233243 | During Preparation for Use | Yes | 58
216601 | The laser system labeling or instructions are incorrect | Rare Occurrence | N/A | Device use information and labeling | Labeling issues | Non-Reportable | Closed | Yes |  | POLAND (POL) | 100340 | Today I received a report over the phone from the FSE that the customer in Poland received the wrong laser.Attached are the delivery vouchers for PLS 100340 for Poland and PLS 100341 for Zwickau. The PLS 100340 was shipped instead to Zwickau and the PLS 100341 was shipped to Poland.Atttached is a photo taken by our FSE during installation in Poland with PLS sn 100341. After further investigation, the labels placed by factory on the crate did not match with the laser inside. | FSE reported wrong SN arrived at customer account. Unknown cause, system was installed without issue. | No Patient Present | N/A | During Installation | Yes | 93
216602 |  |  | N/A |  |  |  | Closed | No |  | POLAND (POL) | 100340 | Today I received a report over the phone from the FSE that the customer in Poland received the wrong laser.Attached are the delivery vouchers for PLS 100340 for Poland and PLS 100341 for Zwickau. The PLS 100340 was shipped instead to Zwickau and the PLS 100341 was shipped to Poland.Atttached is a photo taken by our FSE during installation in Poland with PLS sn 100341. After further investigation, the labels placed by factory on the crate did not match with the laser inside. |  | No Patient Present | N/A | During Installation |  | 0
216655 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100404 | During installation Pin broken on the fiber coupler. | During installation the FSE noted the A34 fiber coupler was damaged on arrival, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20233092 | During Installation | Yes | 62
216666 | Unknown fault code displayed | Unable to replicate complaint | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | IL | UNITED STATES (USA) | 100327 | Unknown fault occurred after calibration completed successfully | FSE was not able to duplicate any faults, system calibrated and operated properly on arrival. | No Injury | 20232327 | During Preparation for Use | Yes | 58
216776 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | MN | UNITED STATES (USA) | 100316 | A broken fiber coupler I found during field action 20232981. | During service the FSE noted a damaged A34 fiber coupler, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20232981 | Routine Service/Not during a case | Yes | 61
216985 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100195 | E106 when trying to calibrate 2 different catheters. Case completed with CVX laser | FSE completed software update and completed calibrations, system is operating properly. | No Injury | 20233281 | During Preparation for Use | Yes | 59 | 106
216998 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100149 | E106. Unable to calibrate | FSE updated the software and completed calibrations, system is operating properly. | No Injury | 20230001 | During Preparation for Use | Yes | 59 | 106
217116 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | IN | UNITED STATES (USA) | 100235 | Rep called to report the system would not calibrate a 1.4 or 1.7 device, reference catheter would calibrate. Case was completed by other means at the same facility and time. | FSE was not able to duplicate the reported problem. Completed calibrations and system is operating properly. | No Injury | 20233285 | During Preparation for Use | Yes | 56
217215 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100101 | Rep called with an Error 106, case was at the end and rep stated physician was done when the failure happened. | FSE was not able to duplicate the reported problem, adjusted alignment and completed calibrations, system is operating properly. | No Injury | 20233288 | After the Procedure | Yes | 57 | 106
217296 | Multiple Fault Codes Displayed | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100259 | Rep reported during calibration the system display is flashing pm due in 2 weeks along with error codes 303 & 310. Case was completed by other means at same facility and time. | FSE found low max energy, system was deinstalled and replaced with SN 100407. | No Injury | 20233299 | During Preparation for Use | Yes | 54 | 310
217305 | The laser system has malfunctioning components | Catheter Coupler Assy | Fault 200 | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100298 | Gerardo Valencia TM called that he couldn’t insert a catheter. Appears one of the coupling switch blades is blocking a switch.Treated patient using CVX300P onsite during same setting. | FSE verified a bent A34 fiber coupler actuator, adjusted. Completed calibrations and system is operating properly. | No Injury | 20233300 | During Preparation for Use | Yes | 54 | 200
217325 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100203 | Rep called to report an E106 after 2 other successful cases. Final case was completed by balloon at same setting. | FSE found rail out of alignment, aligned. Completed calibrations and system is operating properly. | No Injury | 20233312 | During Use in Patient | Yes | 54 | 106
217391 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 230 | Fault 230 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100008 | Error 230 | During service the FSE noted a E230, replaced the high voltage power supply. Completed calibrations and system is operating properly. | No Patient Present | 20233015-BV | Routine Service/Not during a case | Yes | 55 | 230
217394 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100376 | Hannah Edwards called that laser 100376 had a 108 error code.Power cycled the laser to clear error and completed the case. | FSE updated the software and completed calibrations, system is operating properly. | No Injury | 20233313 | During Use in Patient | Yes | 69 | 108
217416 | Clinical catheter fails to calibrate | Fault 108 | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100264 | Rob called to report the system will not calibrate, case was completed with CVX on site at the same setting. | FSE updated the software and completed calibrations, system is operating properly. | No Injury | 20233315 | During Preparation for Use | Yes | 74
217419 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100408 | Rep states the system will not calibrate .9 catheters reliably, he was able to complete calibration with one catheter and was able to complete the case without further issue. | FSE was not able to duplicate, noted the attenuator offset and alignment required adjustment, adjusted. Completed calibrations and system is operating properly. | No Injury | 20233316 | During Preparation for Use | Yes | 74
217499 | Multiple Fault Codes Displayed | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Lower Saxony | GERMANY (DEU) | 100019 | This is a complaint what I noticed during a Service and has nothing to do with a patient case. The customer didn't report any fault. The downloaded LOG files stating the FC310, which is reporable, incombination with FC250. I discussed this fault with Alasdar Mac Cana and as the fault happened in May and from then not anymore we keep monitoring the system and replace the Footswitch if this happens again | During routine service the FSE noted an E310, not duplicated. Completed calibrations and system is operating properly. | No Patient Present | 20230555-BV | Routine Service/Not during a case | Yes | 82 | 310
217514 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100117 | Kevin said they got a Fault code 106. Case was completed with Balloon and Stent. Patient is doing well. | FSE was not able to duplicate the reported problem, replaced the attenuator as a precaution. Completed calibrations and system is operating properly. | No Injury | 20233327 | During Preparation for Use | Yes | 82 | 106
218095 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100117 | Kevin said they got a 106 error code on their laser this morning , waiting to hear back on how the patient was treated and how they are doing.Update from Kevin: Patient is fine, got treatment at 60/60 before the 106 at 60/80. | Multiple E106 after software update, system was deinstalled and replaced with SN 100427. | No Injury | 20233383 | During Use in Patient | Yes | 73 | 106
218194 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Operator Error | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100149 | Dave got an email stating the PLS gave a 201 error when they inserted a 2.5 turbo elite catheter, customer tried another 2.5 Turbo elite catheter and it also gave them the 201 error. Case was finished with a .9 turbo elite , patient is doing well. | Determined customer was using a 2.5mm TE catheter and is not supported for use on the PLS. System is operating properly. | No Injury | 20233400 | During Preparation for Use | Yes | 78 | 201
218241 |  |  | N/A |  |  |  | Closed | No | FL | UNITED STATES (USA) | 100201 | Tony Daley's email said the doctor has had a case where the plaque embolized and occluded the tibial vessels. The patient had to be sent to the hospital and file an incident report. He states that he has never had this issue in the past until the most recent service on the PLS. The laser worked without any errors and they completed the case with it.New Update 11/09/2023:Bri Cucci said :I just wanted to give you an update on the incident from yesterday. I was not in the case, but my partner Jordan was, and he originally informed me that the patient was sent to the hospital. Upon following up with Dr. Sanguily today, the patient was just given an extra dose of Xarelto and was sent home to recover and never went to the hospital. |  | Other | 20233407 | During Use in Patient |  | 71
218279 | The laser system labeling or instructions are not understood by the user | Operator Error | N/A | Device use information and labeling | Instructions For Use (IFU) or Service Manual issues | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100207 | Nature of Complaint: "Missing and confusing technical requirements within the P018732-00 Philips Laser System Operator's Manual, Italian (page 31, section “Precauzioni per la sicurezza”, clause 4). The English equivalent doc is “P019097-00 Philips Laser System Operator's Manual, English” (page 31, section “Safety precautions”, clause 4).The Customer needs to assess adequate eye protection level starting from technical data provided with the equipment. Unfortunately, the Customer could not assess whether the PPE provided was adequate as the information is vague: “Protective eyewear of 5 or greater […]”. What is “5” referring to?Installation was suspended. PPE provided has the following technical specifications:180-315 D LB10 + IR LB4 + M LB6Y (OD10+)The Customer raised concerns due to the unclarity of documentation: they feared that “5” was referring to the LB rating as it is a European requirement following EN207 (standard for full attenuation laser safety eyewear). Instead, it turned out that the Manual refers to OD (Optical Density).The Customer has required a Manual in line with EU requirements (e.g., LB as per EN207) hence with the correct technical data to assess the adequacy of PPE correctly."Actions to resolve (Q15_1):CAS IGT-D provided the Customer with the correct technical requirements the Manual needs to include: “5” refers to OD. Italy H&S Team investigated the issue with the support of Global H&S, and International IGT-D experts from Benelux and Colorado Springs.Resolution (Q94_1):Quality team was informed by GCHO that CFM Galieo do not process Product Feedbacks related to Business units like IGT-Devices, Sleep and Respiratory Care and we also not process Non-Product Feedbacks . Therefore, the record ID 638240831472102222’ is not created and entered into QCM TW. Based on this update, this CF was forwarded to Compaint Expert in IIG who will take care of this CF.Justification of closure (Q93_1):On the basis of the reported information, Quality department established that the customer feedback can be closed here and it will be handled in different way by the CH Market Expert. | Non dispatch phone fix, customer misunderstood the OD rating on the safety glasses thinking it was referring to LB rating. Manual was provided stating Optical Density | No Patient Present | N/A | Not during a case | Yes | 77
218293 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | New Siddhapudur | INDIA (IND) | 100362 | Fault Code:106 with 0.9mm catheter at 70x70 during the clinical procedure.Customer restarted system, they were able to laze at 60x60 but noticed fault code:106 at 70x70.Procedure continued after restart and finished at 60x60. | FSE verified E106, this is post SW update. Found attenuator was not stable, replaced. Completed calibrations and system is operating properly. | No Injury | 20233408 | During Use in Patient | Yes | 77 | 106
218314 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100395 | Hannah Edwards TM called to say PLS100395 will not calibrate .9mm catheters.Able to calibrate reference catheter, 1.4 and 2.7 catheters without issues.Doctor choose to use 1.4mm catheter to complete the case. Patient is fine. | FSE found optical rail out of alignment, aligned. completed calibrations and system is operating properly. | No Injury | 20233410 | During Preparation for Use | Yes | 82
218534 | The reference catheter fails to calibrate | Bad Reference Catheter | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MI | UNITED STATES (USA) | 100120 | Adam Bitner emailed : Customers Ref. Catheter will not calibrate. Clinical Rep’s 0.9 Catheter was tested and calibrated correctly. They would like a tech to come out to look for potential issues. Waiting on Adam to get back to me on how the patient was treated ? update the patient was treated with ballon and Stent and is doing well. | Billable customer declined service, sales provided a reference catheter and system is operating properly. | No Injury | 20233422 | During Preparation for Use | Yes | 78
218539 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100427 | Ben Elkins said Failure On Install, fiber coupler board A34 damaged during shipping and/or assembly.no patient involved | During installation the FSE noted a damaged A34 fiber coupler, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20233402 | Not during a case | Yes | 78
218650 | Clinical catheter fails to calibrate | Catheter Coupler Assy | N/A | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100264 | Pulse Vascular called me to report the laser is having issues calibrating 0.9 and 1.4 catheters again. They receive error calibration failed when trying to calibrate these types of catheters. After multiple attempts, the catheters calibrate ok. They decided to use their back-up laser and would like the PLS laser to be checked again. They were able to complete the cases with the PLS, but decided to use their old CVX laser until the PLS gets checked again. | FSE found a defective A34 fiber coupler, replaced. Completed calibrations and system is operating properly. | No Injury | 20233430 | During Preparation for Use | Yes | 84
218763 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100156 | Rep called to report an E106 during the procedure, case was completed with the other laser system on site during the same setting. | FSE verified E106 and completed software update, completed calibrations and system is operating properly. | No Injury | 20233437 | During Use in Patient | Yes | 78 | 106
218846 | The footswitch connection is not functioning | Damaged Footswitch | Fault 250 | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100285 | Customer called to report the footswitch is not registering front or back unless the cable is pushed in. Case was completed by other means at the same facility and time. | Replacement footswitch was sent to the customer, FSE completed follow up on proper operation. | No Injury | 20233451 | During Preparation for Use | Yes | 77 | 250
218922 | Clinical catheter fails to calibrate | Catheter Coupler Assy | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MI | UNITED STATES (USA) | 100413 | Customer 1.7 catheter will not calibrate. (2) 1.7 catheters were tried and both failed calibration. Reference catheter calibrated properly. Case was completed with CVX laser on-site during the same setting. | FSE verified a defective A34 fiber coupler, replaced and adjusted the front cover. Completed calibrations and system is operating properly. | No Injury | 20233456 | During Preparation for Use | Yes | 77
218944 | The laser system covers are damaged | Interlocks Intermittently Dis-Engage | N/A | Hardware | Enclosure - Structure cracks/breaks/gaps | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | Rep reported the front cover is shifted and system interlocks are tripped. Rep got biomed involved, loosened screws and seated panel. System is operational but Dr wants it checked. Noticed prior to the case and patient will be rescheduled. | FSE verified the front cover shift, rep repositioned. FSE completed follow up, calibrations and system is operating properly. | No Patient Present | 20233460 | Not during a case | Yes | 77
219000 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Optical Rail Mis-Aligned | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100408 | Rep called to report an E103, case was completed with the CVX onsite during the same setting. | FSE found focus lens retainer loose causing misalignment. Tightened and aligned, completed calibrations and system is operating properly. | No Injury | 20233468 | During Use in Patient | Yes | 80 | 103
219018 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OH | UNITED STATES (USA) | 100314 | Rep states having multiple E106 failures, case was completed by other means at the same setting. | FSE verified E106 and completed SW update. Completed calibrations and system is operating properly. | No Injury | N/A | During Preparation for Use | Yes | 80 | 106
219119 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100420 | E106 during the case. Laser was reset, recalibrated the catheter and completed the case with the laser without delay at the same facility. | FSE verified the E106, completed attenuator zero routine which changed the offset. Completed calibrations and system is operating properly. | No Injury | 20233614 | During Use in Patient | Yes | 78 | 106
219455 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | Alabama | UNITED STATES (USA) | 100417 | Fiber coupler damage found on install | During installation the FSE found the A34 fiber coupler damaged, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20233385 | Not during a case | Yes | 71
219469 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 105 | Fault 105 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | QLD | AUSTRALIA (AUS) | 100381 | During an extraction error code 105 appeared on the screen and also noted the system was high in temperature. The physician was extracting the RV lead with roughly 30 seconds of laser activation and had advanced to the SVC, then started to use the outer sheath which advanced to the RV coil while the inner laser sheath remained at the SVC. After a period of time where the physician was attempting to use manual manipulation with the outer sheath to remove the lead, the laser console reverted to standby mode but was then reactivated to ready, how-ever roughly 20sec after being reactivated the physician went to activate the laser when the error code 105 ap-peared with the system note of the high temperature. Unable to continue at this point the physician removed the 16F GlideLight and the laser catheter was disconnect-ed and the PLS was turned off. Turning the PLS back on and recalibrating the catheter the error code did not re-turn and we were able to continue with the procedure. However, the physician was not happy about needing to stop at a crucial stage of the extraction process. The leads were successfully removed and there was no adverse event with the patient. | FSE was not able to duplicate the reported problem, system powered up and operated properly on arrival. | No Injury | 20233641 | During Use in Patient | Yes | 71 | 105
219562 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 231 | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Arizona | UNITED STATES (USA) | 100349 | Laser gave error code 231. Customer was able to clear error and complete the case with laser | FSE was not able to duplicate the E231, did note the A34 fiber coupler is damaged, replaced. Completed calibrations and system is operating properly. | No Injury | 20233658 | During Preparation for Use | Yes | 70 | 231
219586 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 108 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | FRANCE (FRA) | 100066 | during the preparation of the case , technician has constated a crash of the software of PLS .Calibration has been made with the laser prob test and after with the Elca 0,9 for the procedure. She has began the procedure but suddenty she has noticed a message "error 108" on the screen following with a message "call customer service" .and then the procedure was aborted.She has turned off one time and she has remade the operation with a new laser probe and the problem has appeared again.ADDITIONAL INFORMATION OBTAINED 29NOV2023:1. Did the patient have to be transported to another facility to receive treatment? How was case eventually completed? - No2. Did the facility and the physician have alternate options to treat the patient, in case the patient needed them? - yes, stent3. Did aborting the case place the patient at an increased risk? - no increased risk for the patient. All finished well4. Was the patient open and access obtain? - yes good results5. Was the patient already under anesthesia or sedation? - local anesthesia6. Was angiogram performed? - n/aWhat is meant by “good results”?For example, did patient receive a stent or any other treatment to complete the procedure that day?- Yes effectively , the lesion off the patient has been prepared with the laser prob in order to put a stent in the middle LCA The result was perfect with no injury for the patient.The angiogram has been performed and looks good but with the French protection law I can’t give you without an agreement.Did the catheter effectively go through the lesion before the PLS gave the fault code?-To what I understood , the first catheter could not cross the lesion du to the problem from the PLS and the default.That’s why they have opened a second for treat the patient. The customer thought that the problem came from the laser probeAnd with the second catheter , they have succeed the procedure despite the errors still presents | FSE verified and completed SW update, completed calibrations and system is operating properly. | No Injury | 20230635-BV | During Preparation for Use | Yes | 69 | 108
219692 | The laser system will not turn on | Manufacturing Error | N/A | Power | Will not turn on | Non-Reportable | Closed | Yes | KS | UNITED STATES (USA) | 100508 | During installation laser would not turn on. Install report CRN20233620 | During installation the FSE noted the system would not fire on power up. FSE found the fiber cable disconnected, reconnected and was able to complete the installation. | No Patient Present | 20233620 | During Installation | Yes | 75
219693 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 231 | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100346 | At the end of the case error 231 occurred | FSE found a loose connection to the HVPS, tightened. Completed SW update and calibrations, system is operating properly. | No Injury | 20233676 | After the Procedure | Yes | 71 | 231
219792 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100148 | E106 when trying to calibrate. CVX laser on site was used to complete this case | FSE was not able to duplicate the reported problem, completed the SW update and calibrations. System is operating properly. | No Injury | 20233680 | During Preparation for Use | Yes | 71 | 106
219796 | Unknown fault code displayed | Unable to replicate complaint | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100269 | Laser not working, unsure of fault numberUPDATE 05DEC2023:catheter is being returned for failure analysis per request of tech support | FSE was able to verify the failure to calibrate but system is within specification and calibrates all service catheters. This system has SW update completed. | No Injury | 20233681 | During Preparation for Use | Yes | 83
219910 | The laser system generates a buzzing noise and won't turn on | E-Stop Engaged During Power-Up | Unknown | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100205 | When turning laser on at start of day the laser gives a constant buzzing sound. No patient involvement | FSE found the E-stop switch engaged during power up causing the buzzing sound, released the switch and system powered up normally. | No Patient Present | 20233690 | Not during a case | Yes | 69
219912 | Catheter coupler issue | Catheter Coupler Assy | Unknown | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | MI | UNITED STATES (USA) | 100413 | “unsuccessful Calibration” with 1.7 and 2.0 catheters. | FSE found the front cover making contact with the A34 fiber coupler, replaced and adjusted door, completed calibrations and system is operating properly. | No Injury | 20233691 | During Preparation for Use | Yes | 70
220002 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100094 | Ken- Account got an unable to calibrate, while trying to calibrate the first clinical catheter. They tried a second Elca catheter and they were able to calibrate and complete the case with the laser. patient is doing well. | FSE found the beam waist misaligned, adjusted and completed calibrations. System is operating properly. | No Injury | 20233701 | During Preparation for Use | Yes | 71
220038 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100332 | Stephanie saidIt was error code 106. I powered down the laser after it prompted (as it does not go away any other method) and then turned it back on and recalibrated it. It then worked for a bit, until it occurred again and I repeated that process. Case was completed with laser and patient is doing well. | FSE verified and updated SW, completed calibrations and system is operating properly. | No Injury | 20233705 | During Use in Patient | Yes | 79 | 106
220128 | Unknown fault code displayed | Electro / Mechanical | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Florida | UNITED STATES (USA) | 100250 | Don Slate called to say the Laser gave a fault but he did not remember what it was with a 1.4 clinical catheter during the case, They removed the 1.4 catheter and recalibrated and was able to complete the case with no further issues. He also thought it was taking longer than normal for the energy to reach the needed level for calibrating the 1.4 cath.it did calibrate. Patient is doing well. | FSE found the optical rail misaligned causing issues with coupling, cleaned optics and aligned. Completed calibrations and system is operating properly. | No Injury | 20233708 | During Use in Patient | Yes | 78
220144 | The footswitch cable or cover is damaged | Damaged Footswitch | Unknown | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100158 | Ryan DeLucas TM: Customer is in need of a service call due to a broken pedal. laser was pushed into a wall, No cases involved.Trying to find out if this is just a Footswitch needed or something else. If it is just a footswitch I will change this to a non dispatched call and we can ship them a footswitch.*Update this is just a footswitch so please ship a footswitch to account, waiting on Andy to give me the contact info to receive the footswitch | Replacement footswitch was sent to the customer. FSE completed follow up on receipt and proper operation. | No Patient Present | 20233710 | Not during a case | Yes | 100
220237 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100357 | Wade said they got a 106 error code , swapped catheters and recalibrated and was able to complete the case. Patient is doing well. | FSE found the beam waist to be small and out of alignment, adjusted. Completed calibrations and system is operating properly. | No Injury | 20233722 | During Preparation for Use | Yes | 77 | 106
220343 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | SC | UNITED STATES (USA) | 100377 | Sales Rep Nicole McKee @ 843-906-1451 , After the cases were completed today the laser beeped and had an Overtemp message on the screen. Case went fine and patient is doing well. | FSE was not able to duplicate the reported problem, rerouted and reseated inline fuses. Completed calibrations and system is operating properly. | No Injury | 20233735 | After the Procedure | Yes | 76
220359 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100269 | Umberto said they are getting an unable to calibrate with the 1.4 clinical catheters x2. Patient was treated with angioplasty, and is doing fine. The 1.4 today the lot number isFAY23H08A | FSE found reported error was 310, not duplicated. System was deinstalled and replaced due to customer confidence. | No Injury | 20233741 | During Preparation for Use | Yes | 84
220365 | The laser system generates unusual noise | Unable to replicate complaint | N/A | Other | Other | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100090 | Juan Sales @ 915-603-8417 Clinical rep said the PLS laser is making an unusual noise while trying to calibrate a catheter. The laser calibrates the catheter and is working fine just concerned about the unusual noise. Case was completed as planned with laser and patient is doing well. | FSE was not able to duplicate any unusual noises, system powered up and operated properly on arrival. | No Injury | 20233742 | During Preparation for Use | Yes | 77
220440 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100372 | Mike said they got a 103 error while trying to calibrate the catheter, They reset power and recalibrated with a new catheter and was able to complete the case with laser. Patient is doing well. | FSE found a damaged fiber coupler PCB, replaced. Completed calibrations and system is operating properly. | No Injury | 20233746 | During Preparation for Use | Yes | 76 | 103
220522 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100410 | E600 during the case. Case was completed with the laser without delay at the same facility. | FSE verified and reseated connection to the fuses and PCN. Completed calibrations and system is operating properly. | No Injury | 20233752 | During Use in Patient | Yes | 72 | 600
220695 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) | Electro / Mechanical | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | 30 Day Reportable | Closed | Yes | OR | UNITED STATES (USA) | 100198 | E600 during the case. FSE to find out how the case was completed.Update from Eleazar Sanchez 28DEC2023Patient was on the table, it was a planned diagnostic procedure with possibility of intervention. Conscious Sedation was on board. During the diagnostic portion of the procedure patients distal tibial arteries were found to be occluded. This was found to be inappropriate cause of patience pain and discomfort and further symptoms. Opening of the arteries was necessary in order to relieve symptoms. With the failure of the laser for the procedure was aborted and planned for a future time once the laser was up and going.Laser errored out prior to opening arteries.Spoke to rep. David Bartlemay on 09JAN2024Since the Dr. wasn't able to cross the distal tibilas due to occlusion and E600 with the PLS there was no other option to cross. | FSE found loose connection to PCN, reconnected. Completed calibrations and system is operating properly. System repairs were completed on Dec 14, 2023. | No Injury | 20233768 | During Use in Patient | Yes | 84 | 600
220722 |  |  | Fault 108 |  |  |  | Closed | No |  | FRANCE (FRA) | 100066 | during the preparation of the case , technician has constated a crash of the software of PLS .Calibration has been made with the laser prob test and after with the Elca 0,9 for the procedure. She has began the procedure but suddenty she has noticed a message "error 108" on the screen following with a message "call customer service" . |  | No Injury | N/A | During Preparation for Use |  | 0 | 108
220723 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 231 | Fault 108 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | FRANCE (FRA) | 100066 | during the preparation of the case , technician has constated a crash of the software of PLS .Calibration has been made with the laser prob test and after with the Elca 0,9 for the procedure. She has began the procedure but suddenty she has noticed a message "error 108" on the screen following with a message "call customer service" .ADDITIONAL INFORMATION OBTAINED 13DEC2023:Did the facility and the physician have alternate options to treat the patient, in case the patient needed them? Yes , they have another option with Rotapro and stent Did aborting the case place the patient at an increased risk? Was this an emergent case?No increased risk for the patient and it wasn’t an emergency caseWas the patient open and access obtained?The patient wasn’t open and no access obtainedWas the patient already under anesthesia or sedation?The patient was already under anesthesiaWas this a LM, Coronary, or peripheral case?Coronary casePer Physician Assessment 13DEC2023, this event determined to be non reportable. | Problem was actually an E231, not duplicated. Completed calibrations and system is operating properly. | No Injury | 20230788-BV | During Preparation for Use | Yes | 72 | 108
220729 | Unknown fault code displayed | Unknown Fault | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Tokyo | JAPAN (JPN) | 100331 | PLS displayed error messageThe customer saw error message. The customer restarted the system to clear the error message. After reboot the system. The PLS worked properly. Was the Device in Clinical Use at the Time the Issue was discovered? ? NoWas any Patient or User Harmed? ?NoThere was no impact to patient. | FSE states that customer refused services, system is operating properly and error has not returned. | No Injury | 20233776 | During Preparation for Use | Yes | 71
220767 | Clinical catheter fails to calibrate | Fault 310 | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100269 | Failure to calibrate. Laser was reset and calibrated the catheter properly. Case was completed with the laser without delay at the same facility. | FSE determined the error was a 310, system was deinstalled and replaced due to customer confidence and multiple failures. | No Injury | 20233770 | During Preparation for Use | Yes | 71
220794 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100298 | E106 with a 0.9mm catheter at higher rep rates. Case was completed with their spare CVX without delay at the same facility. | FSE found a noisy PES, replaced, also adjusted the optical rail alignment. completed calibrations and sysetm is operating properly. | No Injury | 20233772 | During Use in Patient | Yes | 76 | 106
220796 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100351 | E106 with a 1.4mm catheter when increasing the settings to 60/80. Case was not completed with the laser. It was completed with a balloon without delay at the same facility. | FSE verified and completed software update, completed calibrations and system is operating properly. | No Injury | 20233773 | During Use in Patient | Yes | 78 | 106
220874 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 231 | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100106 | Rep called to report an E231 at the end of the procedure. The case was completed with no harm to the patient. | FSE was not able to duplicate the E231, completed calibrations and system is operating properly. | No Injury | 20233778 | After the Procedure | Yes | 71 | 231
220885 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100394 | Rep called to report an Error 500, case was completed by mechanical device during the same setting. | FSE found a defective shutter, replaced. Completed calibrations and system is operating properly. | No Injury | 20233780 | During Preparation for Use | Yes | 70 | 500
221026 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | N/A | ITALY (ITA) | 100109 | Error code 500-501 - System error (SELF_TEST_FAILED) | FSE found a defective PES sensor, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20230785-BV | Not during a case | Yes | 77 | 500
221083 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100414 | Rep called to report an error 221 when the PLS powers up that clears after a restart. Case was completed with the laser. | FSE found a defective attenuator encoder motor, replaced. Completed calibrations and system is operating properly. | No Injury | 20233936 | During Preparation for Use | Yes | 77 | 221
221088 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Az | UNITED STATES (USA) | 100358 | Rep called to report an E106 during the procedure, rebooted and calibrated fine but gave another E106 after 2 passes with a .9 device 60/60 setting. Case was completed by other means during the same setting. | FSE verified and completed SW update, completed calibrations and system is operating properly. | No Injury | 20233938 | During Use in Patient | Yes | 87 | 106
221172 | The laser system has broken components | Damaged Laser System Cover | N/A | Hardware | Accessory issues | Non-Reportable | Closed | Yes | Noord Brabant | NETHERLANDS (NLD) | 100391 | The system arrived in Bomi warehouse from COS with damage to its rear handle. The rear handle was replaced during System upgrade prior to shipping to customer for installation . Handle damaged and replaced on 15 december 2023 resolved dispatch | During installation the FSE noted damaged paint on rear handle, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20230669-BV | Not during a case | Yes | 76
221176 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Florida | UNITED STATES (USA) | 100250 | Rep called to report an E106, case was completed with the CVX on site during the same setting. | FSE verified error, completed atten zero and optical rail alignments and calibrations, system is operating properly. | No Injury | 20233944 | During Use in Patient | Yes | 76 | 106
221198 | The laser system does not provide reliable performance | Unable to replicate complaint | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100026 | Customer called to report an unrecorded error during use. Said it took several times to calibrate and failed during use. They used the CVX system onsite to complete the case in the same setting. | FSE verified failure to calibrate, found optical rail out of alignment, aligned. Completed calibrations and system is operating properly. | No Injury | 20233946 | During Use in Patient | Yes | 77
221371 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100026 | Rep called to report the system failed to calibrate, case was completed with the CVX on site during the same setting. | FSE was not able to duplicate the failure to calibrate, completed atten zero and alignments, completed calibrations and system is operating properly. | No Injury | 20233985 | During Preparation for Use | Yes | 74
221479 | Unknown fault code displayed | Unknown Fault | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AR | UNITED STATES (USA) | 100196 | Rep called to report an unknown error on power up, customer cycled power and error cleared, was able to complete the case. | Customer refused services and is using the system without issue. | No Injury | 20234004 | During Preparation for Use | Yes | 80
221490 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100160 | Rep called to report an E106, case was completed by other mean during the same setting | FSE verified E106, completed SW update and calibrations. System is operating properly. | No Injury | 20234005 | During Use in Patient | Yes | 80 | 106
221651 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100519 | David Fields 914-772-3713 said the account got a 103 fault code they were able to reset and recalibrate the clinical catheter and complete the case. Patient is doing well. | FSE found damaged fiber coupler PCA, replaced. completed calibrations and system is operating properly. | No Injury | 20234013 | During Preparation for Use | Yes | 75 | 103
221778 | The laser system is damaged during transport | Catheter Coupler Assy | N/A | Hardware | Accessory issues | Non-Reportable | Closed | Yes | Florida | UNITED STATES (USA) | 100526 | The Fiber Coupler Assy was damaged on install / Arrival. | FSE found a damaged fiber coupler PCA on installation, replaced. Completed calibrations and system is operating propelry. | No Patient Present | 20233973 | During Installation | Yes | 73
221871 | The laser system generates unusual noise | Electro / Mechanical | N/A | Other | Other | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100060 | Constant buzzing sound when laser is plugged in | FSE found a defective interlock switch, replaced. Completed calibrations and system is operating properly. | No Injury | 20240016 | During Preparation for Use | Yes | 71
221937 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | Error 221 when trying to use the laser. Case completed by other means | FSE found a defective attenuator encoder, replaced. Completed calibrations and system is operating properly. | No Injury | 20240020 | During Preparation for Use | Yes | 64 | 221
222149 | The laser system generates unusual noise | E-Stop Engaged During Power-Up | N/A | Other | Other | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100205 | when setting the circuit breaker to the 'On" position, a constant buzzing sound is heard. The sales rep reset the EMO button with the same result. Case was not completed with this laser. The case was completed with their CVX300 without delay at the same facility. | FSE found EMO depressed during start up, released and system powers up and operates properly. | No Injury | 20230034 | During Preparation for Use | Yes | 73
222263 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100330 | E103 during calibration. Sales Rep to report how the case was completed and condition of the patient.Update: Cather was recalibrated and the case was completed with the laser without delay at the same facility. | FSE was not able to duplicate the reported problem, completed SW update and calibrations, system is operating properly. | No Injury | 20240044 | During Preparation for Use | Yes | 77 | 103
222477 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100122 | Rep called to report multiple E106/108's over the past 4 months, DMR review indicates these were not reported, last routine service was September 2023 with no issues. Unknown how any cases which experienced E106 were completed and how. Rep is requesting SW update. | FSE verified and completed SW update and calibrations. System is operating properly. | No Patient Present | 20240056 | Not during a case | Yes | 71 | 106
222488 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 303 | Fault 303 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Alabama | UNITED STATES (USA) | 100178 | E303 during a procedure. Laser was restarted and the case was completed with the laser without delay at the same facility. | FSE found optical rail out of alignment and lower energy output. Aligned and completed calibrations and system is operating properly. | No Injury | 20240066 | During Use in Patient | Yes | 71 | 303
222495 | Multiple Fault Codes Displayed | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | WA | UNITED STATES (USA) | 100119 | E103 & 108. Unknown how the case was completed as they were in the middle of the case.12Jan2024 Update: Case was not completed with the laser. It was completed with a balloon without delay at the same facility. | FSE verified and completed SW update, completed calibrations and system is operating properly. | No Injury | 20240070 | During Use in Patient | Yes | 71 | 103
222580 | The power cord cable is damaged | Damaged Power Cord | N/A | Power | Will not turn on | Non-Reportable | Closed | Yes | Georgia | UNITED STATES (USA) | 100233 | Laser intermittently shutoff. Biomed found that the power cord plug is damaged and possible damage to the laser receptacle. Case was completed with the laser without delay at the same facility. | FSE verified damaged power cord and replaced, completed calibrations and system is operating properly. | No Injury | 202400784 | During Preparation for Use | Yes | 70
222583 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | Unknown | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100540 | Laser failed to calibrate two 1.4mm catheters. Sales Rep did not state how the case was completed.Update 12Jan2024: Case was completed with their backup CVX laser without delay at the same facility. | FSE found the optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20240077 | During Preparation for Use | Yes | 70
222620 |  |  | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | Ishikawa | JAPAN (JPN) | 100387 | Fault 231 occurred during preparation for use. The customer power cycled the system and it operated properly. FSE asked the customer to wait and see.Was the Device in Clinical Use at the Time the Issue was discovered? -> NoWas any Patient or User Harmed? -> NoAfter power cycling, the system operated properly.This complaint was closed by phone fix. FSE was not dispatched and no parts were sent to the customer site. |  | No Patient Present | NA | During Preparation for Use |  | 59 | 231
222621 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 231 | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Ishikawa | JAPAN (JPN) | 100387 | Fault 231 occurred during preparation for use. The customer power cycled the system and it operated properly. FSE asked the customer to wait and see.Was the Device in Clinical Use at the Time the Issue was discovered? -> NoWas any Patient or User Harmed? -> NoAfter power cycling, the system operated properly.This complaint was closed by phone fix. FSE was not dispatched and no parts were sent to the customer site. | FSE instructed customer to power cycle the system and error cleared, system is operational. Non dispatch phone fix. | No Patient Present | N/A | During Preparation for Use | Yes | 59 | 231
222668 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100541 | Rep called to report a failure to calibrate with a 16F device. Case was completed using a 14F and upsize to a 13F TightRail. | FSE found the optical rail misaligned, aligned and completed calibrations, system is operating properly. | No Injury | 20240084 | During Preparation for Use | Yes | 66
222741 | Multiple Fault Codes Displayed | Fault 221 | Fault 301 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100536 | Problem: Error 301 the first time I tried to fire the laser. Intermittent Error 221 during subsequent reboots.Cause: Faulty attenuator assembly. (P/N: P010375)Solution: Replaced attenuator assembly and ran attenuator Zero procedure. No further repeats of Error 301 or 221. | During installation the FSE noted a E301 and E221, replaced the encoder motor, completed calibrations and system is operating properly. | No Patient Present | 20233991 | During Installation | Yes | 66 | 301
222743 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100142 | Rep called to report failure to calibrate using 2 different .9 devices. Case was completed by other means during the same setting. | FSE found the optical rail misaligned, aligned and completed calibrations, system is operating properly. | No Injury | 20240243 | During Preparation for Use | Yes | 66
222744 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | Az | UNITED STATES (USA) | 100358 | Rep called to report failure to calibrate with a .9 device, case was completed by other means during the same setting. | FSE found the optical rail out of alignment, aligned and completed calibrations, system is operating properly. | No Injury | 20240244 | During Preparation for Use | Yes | 66
222746 | The footswitch cable or cover is damaged | Damaged Power Cord | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | Florida | UNITED STATES (USA) | 100250 | The power cable is cracked exposing the wires. Case was completed with the laser at the same facility. No harm the patient. Ship a new power cable. | Non dispatch, replacement power cord was sent to the customer, FSE completed follow up on receipt and proper operation. | No Injury | 20240248 | After the Procedure | Yes | 66
222765 | Clinical catheter fails to calibrate | Electro / Mechanical | N/A | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100427 | Rep called to report failure to calibrate with a 2.0 TE, power cycled and was able to get it to calibrate and complete the case. Also reported is an unusual burning smell. | FSE found the atten zero was off and also aligned the optical rail. Completed calibrations and system is operating properly. | No Injury | 20240249 | During Preparation for Use | Yes | 58
222985 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100149 | Customer called to report an E106 using a 14F and 16F device, power cycled the system to try and re-calibrate and the system indicated an E221 on power up. Case was completed by other means. SW has been updated. | FSE found a defective attenuator encoder causing 106/221, replaced and completed calibrations, system is operating properly. | No Injury | 20240257 | During Use in Patient | Yes | 64 | 106
223000 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100162 | Rep called to report an E277 per the customer, no patients were involved account testing the system. Expect finding to actually be an E221 as this is happening during power up and there are no E277 listed for the PLS. | FSE found a defective attenuator motor causing an E221, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20240259 | Not during a case | Yes | 64 | 221
223007 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Fl | UNITED STATES (USA) | 100179 | Rep reported an E221 during a case, was completed with a balloon at the same setting. | FSE found a defective attenuator, replaced and completed calibrations, system is operating properly. | No Injury | 20240260 | During Use in Patient | Yes | 64 | 221
223106 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100026 | Customer reported failure to calibrate with a 1.7mm device, case was completed with the CVX system on site using the same catheter. | System has had multiple failure to calibrate issues and will be deinstalled and replaced with SN 100573. | No Injury | 20240266 | During Preparation for Use | Yes | 63
223306 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Gujarat | INDIA (IND) | 100338 | Customer has reported Error 106 after boot-up. | FSE completed SW update and calibrations, system is operating properly. | No Patient Present | 20240283 | Not during a case | Yes | 59 | 106
223318 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100197 | Don Slate : Post successful calibration this 2.0 Turbo Elite experienced a 106 PLS error code while attempting to use in the patient. The issue could not be rectified, and atherectomy was aborted.ADDITIONAL INFORMATION OBTAINED 22JAN2024:1. Did the facility and the physician have alternate options to treat the patient, in case the patient needed them? Yes2. Did aborting the case place the patient at an increased risk? No3. Was the patient open and access obtain? Not clear what you are asking as this was a percutaneous procedure via the common femoral artery. 4. Was the patient already under anesthesia or sedation? Yes, IV Conscious Sedation5. Was angiogram performed? YesADDITIONAL INFORMATION OBTAINED 23JAN2024:Don just called me back and said the case abortion thing was with the laser aborted and the case was completed with balloon and stent and patient is well. The laser worked for over half of the procedure before they got the 106 so the doctor was able to just balloon and stent to complete the procedure. | FSE was not able to verify the reported problem, system calibrated and operated properly. | No Injury | 20240286 | During Use in Patient | Yes | 59 | 106
223429 | Clinical catheter fails to calibrate | Unknown Fault | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100103 | Morgan said they tried to calibrate the catheter and got an unable to calibrate message. so they powered down the laser and tried a new catheter and got the unable to calibrate message again. Waiting on Morgan to call me back about how the patient will be treated as the patient was still on the table.Update from Morgan at the account, the patient was treated with another atherectomy device and is doing well. | FSE was not able to duplicate the failure to calibrate but did note debris on the face of the CES detector, cleaned. Completed calibrations and system is operating properly. | No Injury | 20240288 | During Preparation for Use | Yes | 58
223446 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | AR | UNITED STATES (USA) | 100208 | Josh said the laser was unable to calibrate the .9 catheter Reset power and still got Unable to calibrate while trying to calibrate the .9. Case was complete with a Philips Scoring Balloon. Patient is doing well. Josh Rowley 501-622-0069 | FSE was not able to duplicate the failure to calibrate, completed alignments and calibrations and system is operating properly. | No Injury | 20240293 | During Preparation for Use | Yes | 58
223546 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 240 | Fault 240 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100127 | TM Tyler Luellen @ 636-399-2430- Said the laser got a fault 240 when they powered it up this morning, They reset power to laser and it powered up correctly so they were able to complete 3 cases after they got the 240 fault with no further problems. Patients are all doing well. | FSE confirmed the E240 but was not able to duplicate, completed calibrations and the system is operating properly. | No Injury | 20240297 | During Preparation for Use | Yes | 57 | 240
223552 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 231 | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | ID | UNITED STATES (USA) | 100114 | Philips Rep Cragin Orr @ 719-659-1679 - said they got a 231 fault , while setting up for the case.Reset power to laser and the laser worked with no issues for the case. Patient is doing well. | Customer refused the services, restarted the system and fault cleared. | No Injury | 20240300 | During Preparation for Use | Yes | 57 | 231
223652 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100554 | Jonathan said they were doing an in-service on their newly installed laser and they calibrated the 2.5 ref cath but got an unable to calibrate the 1.7 and a .9 clinical catheter. Unable to calibrate message for both. They were able to complete the in-service with the 2.5 cath. | FSE was not able to duplicate the reported problem, completed atten zero and calibrations, system is operating properly. | No Patient Present | 20240306 | Not during a case | Yes | 62
223658 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | DENMARK (DNK) | 100166 | Customer got an error message 106 after 10 seconds. They have calibrated the system with the “dummy”Systems works. Then they switch to the Turbo Power and after 10 seconds they received this error message 106. | FSE could not duplicate the reported problem, replaced the encoder motor and completed calibrations, system is operating properly. | No Injury | 20240060-BV | During Preparation for Use | Yes | 69 | 106
223672 | The laser system has broken components | Damaged Laser System Cover | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | IL | UNITED STATES (USA) | 100556 | PLS handle broke during shipping. It’s the main handle used to push the laser system. | During installation the FSE noted a damaged rear handle, will be replaced under a separate call per customer request. Completed calibrations and system is operating properly. | No Patient Present | 20240247 | During Installation | Yes | 62
223752 | Clinical catheter fails to calibrate | Electro / Mechanical | Unknown | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100262 | Laser has trouble calibrating and will sometimes shut off during a case. Case was completed with laser. | FSE was not able to duplicate the reported problem, updated the software and replaced attenuator. Completed calibrations and system is operating properly. | No Injury | 20240325 | During Use in Patient | Yes | 73
223843 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100183 | Multiple Error 106 during procedure.Laser was restarted to clear error. Case was completed by Laser.No patient injury. | FSE completed SW update and calibrations, system is operating properly. | No Injury | 20240334 | During Preparation for Use | Yes | 65 | 106
224049 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100315 | Laser will not clibrate | FSE found optical rail was mis aligned, aligned. Completed calibrations and system is operating properly | No Injury | 20240350 | During Preparation for Use | Yes | 56
224135 | The laser system energy delivery cannot be controlled by the user | Unable to replicate complaint | Unknown | Other | Other | Non-Reportable | Closed | Yes | AR | UNITED STATES (USA) | 100304 | Not enough energy was emitted from the laser post laser images. The laser did nothing to the lesion. | FSE did not find any issue with the system, it calibrated and operated properly on arrival. | No Injury | 20240352 | During Use in Patient | Yes | 56
224175 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Software/ Firmware (excluding privacy or security) | Failure to upgrade | Non-Reportable | Closed | Yes | Ma | UNITED STATES (USA) | 100061 | Failed after software upgrade. Received Fault 221. | During routine service and SW update the FSE noted an E221, replaced the encoder motor. Completed calibrations and system is operating properly. | No Patient Present | 20233560 | Routine Service/Not during a case | Yes | 55 | 221
224266 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Optical Rail Mis-Aligned | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100080 | E103 calibrating the reference catheter. Reset laser and calibrated the catheter properly. Case was completed with the laser without delay at the same facility. | FSE found the optical rail out of alignment, aligned. Completed calibrations and system is operating properly. | No Injury | 20240360 | During Preparation for Use | Yes | 52 | 103
224277 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AR | UNITED STATES (USA) | 100553 | E103 with two different 1.7mm Rx catheters. Case was not completed with the laser. it was completed with a Phoenix device without delay at the same facility. | FSE could not duplicate, completed optical alignment and calibrations, system is operating properly. | No Injury | 20240362 | During Preparation for Use | Yes | 52 | 103
224396 | The laser system turns off during a case | Unable to replicate complaint | N/A | Power | Other | Non-Reportable | Closed | Yes | TN | UNITED STATES (USA) | 100539 | Customer accidently powered off during a case, and after powering on was unable to calibrate on two 12fr catheters, but was able to calibrate the reference 2.5mm catheter.Case was completed with the CVX laser without delay at the same facility. | FSE was not able to duplicate the failure to calibrate, system calibrated and operated properly on arrival. | No Injury | 20240375 | During Use in Patient | Yes | 62
224443 | Multiple Fault Codes Displayed | Damaged Footswitch | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MI | UNITED STATES (USA) | 100353 | Attempted to use 1.7 turbo elite. Upon startup the laser displayed a 250 system failure code. Attempted to plug foot pedal in to both ports. Attempted to calibrate and received 310 system failure code. Restarted system and repeated the process with the same results.Laser catheter never entered the patient. No injury to patient. Decided to use CSI 1.5 solid and completed the procedure as planned without delay at the same facility. | FSE found defective footswith and replaced. Completed calibrations and system is operating properly. | No Injury | 20240387 | During Preparation for Use | Yes | 61 | 310
224466 | The laser system has broken components | Catheter Coupler Assy | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 1000571 | Found Fiber Coupler Damaged on install. | During installation the FSE noted a damaged coupler A34 PCB, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20240310 | During Installation | Yes | 62
224586 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100178 | E310 and failed to calibrate Turbo Elite 2.0 during the preparation for the case. Case was not completed with the laser. It was completed with a Phoenix 1.5 without delay at the same facility. | FSE will deinstall and replace the system with SN 100597 due to multiple failures and customer confidence. | No Injury | 20240392 | During Preparation for Use | Yes | 61 | 310
224591 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100549 | Calibration errors using a 1.4, 1.7 & 2.0 catheters. Eventually the catheter calibrated and the case was completed with the laser without delay at the same facility. | FSE was not able to duplicate, adjusted the atten offset and alignments. Completed calibrations and system is operating properly. | No Injury | 20240393 | During Preparation for Use | Yes | 74
224622 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CO | UNITED STATES (USA) | 100167 | Multiple E106 during the case. Case was not completed with the laser. It was completed with angioplasty without delay at the same facility. | FSE completed SW update and calibrations, system is operating properly. | No Injury | 20240398 | During Use in Patient | Yes | 61 | 106
224715 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | LA | UNITED STATES (USA) | 100103 | 1.4mm catheter failed to calibrate. Customer cleaned the detector and tried to recalibrate. It failed again. Case was completed with a balloon without delay at the same facility. | FSE not able to duplicate failure to calibrate, replaced attenuator as a precaution and completed calibrations, system is operating properly. | No Injury | 20240404 | During Preparation for Use | Yes | 73
224721 |  |  | Unknown | Calibration | Unable to Calibrate |  | Closed | No | LA | UNITED STATES (USA) | 100103 | catheter 1.4mm failure to calibrate. |  | No Injury | 20240404 | During Preparation for Use |  | 6
224798 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Tamil Nadu | INDIA (IND) | 100095 | Customer has reported error 106 during procedure. Need to visit site and check. Case was completed, they were testing routine calibration of catheter. | FSE verified and found humidity/dust on the optics and sensors, cleaned. Completed calibrations and system is operating properly. | No Injury | 20240409 | During Preparation for Use | Yes | 65 | 106
224855 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100557 | Rep states failure to calibrate using a 2.0 T.E. device. Customer used the other PLS on site to complete the case using the same catheter. | FSE verified failure to calibrate and found optical rail misaligned, aligned and completed calibrations. System is operating properly. | No Injury | 20240422 | During Preparation for Use | Yes | 51
224977 | The laser system has malfunctioning components | Damaged Laser System Cover | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | JAPAN (JPN) | 100501 | We received PLS SN100501 at the end of last year. And FSE Fukushima-san was verifying the system. And He found two issues on the system. Damaged front cover and Error 600. | FSE noted a damaged front enclosure during installation, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20240051 | Routine Service/Not during a case | No | 50 | 600
224978 | The laser system will not turn on | System Would Not Turn On | N/A | Power | Will not turn on | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100571 | Rep called to report a constant buzzing during power up and system will not turn on. This was found during an in service and there are no patients involved. | FSE found front cover interlock is intermittent, seated front cover. Completed calibrations and system is operating properly. | No Patient Present | 20240932 | Not during a case | Yes | 68
224982 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | HI | UNITED STATES (USA) | 100177 | Rep called to report an E106 and request a SW update. Case was completed at lower setting with the system. | FSE verified E106, found rail optics misaligned, aligned and recalibrated sensors. Completed calibrations and system is operating properly. | No Injury | 20240570 | During Use in Patient | Yes | 63 | 106
225083 | The laser system covers are damaged | Damaged Laser System Cover | N/A | Hardware | Enclosure - Structure cracks/breaks/gaps | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100132 | Rep reported a damaged front enclosure, system is operating properly and case was completed without issue. | FSE verified damaged front enclosure and replaced. completed calibrations and system is operating properly. | No Injury | 20240580 | After the Procedure | Yes | 61
225451 | The laser system has broken components | Emergency Switch | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100024 | During demo during an in-service the emergency button broke off. | FSE verified damaged E stop, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20240597 | Not during a case | Yes | 48
225461 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 211 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | Tokyo | JAPAN (JPN) | 100331 | Error 221 occurred during calibration. The customer powered cycle the system and it worked normally. This problem is intermittent.Was the Device in Clinical Use at the Time the Issue was discovered? -> NoWas any Patient or User Harmed? -> NoFSE will go in-site and investigate the problem. | FSE verified error 221, replaced the attenuator motor, completed calibrations and system is operating properly. | No Patient Present | 20240599 | Not during a case | Yes | 48 | 211
225508 | Clinical catheter fails to calibrate | Optical Rail Mis-Aligned | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100028 | Laser will not calibrate | FSE found optical rail misaligned, aligned. Completed calibrations and system is operating properly. | No Injury | 20240600 | During Preparation for Use | Yes | 61
225534 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100141 | E106 when trying to use the laser. Case completed with 2nd laser on-site | FSE verified E106 and was not able to correct, system was deinstalled and replaced with SN 100614. | No Injury | 20240602 | During Preparation for Use | Yes | 60 | 106
225843 | Clinical catheter fails to calibrate | Unknown Fault | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100540 | Dave Rough TM @ 908-963-7486 , While trying to calibrate the 1.4 clinical catheter they got a will not calibrate message. they used their CVX300 laser to complete the case and patient is doing well. | FSE verified failure to calibrate, completed alignment of optical rail and attenuator zero routine. Completed calibrations and system is operating properly. | No Injury | 20240630 | During Preparation for Use | Yes | 51
225850 | The power cord connector is damaged | Damaged Power Cord | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | MO | UNITED STATES (USA) | 100230 | At the stress release area of the power cord, it is broken. The internal wires are visible and able to be touched. Ship another power cord. Case completed with the laser. No harm to patient.Additional information received 29FE2024:The strain relief on the cord right where it plugs into the laser console was broken, so you could see the colored wires inside. The colored insulation was showing and not bare copper wires. | Replacement power cord was sent to the customer, FSE completed follow up on receipt and proper operation. | No Injury | N/A | During Use in Patient | Yes | 51
225871 | Clinical catheter fails to calibrate | Unknown Fault | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100079 | Jeff said the 12F and 14F clincal catheters gave an Unable to Calibrate message on their PLS laser. They completed the case with their CVX300 laser. Patient is doing well. | FSE was able to duplicate the failure to calibrate, found rail alignment out and beam size, adjusted. Completed calibrations and system is operating properly. | No Injury | 20240642 | During Preparation for Use | Yes | 51
225953 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 201 | Fault 201 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100351 | Philips Rep Jacey Aaron @ 706-669-4995, said they got a Fault code 201 when the powered up the laser. before inserting a catheter. Reset power and it came back on with no error , they were able to calibrate and use the laser for the case. Patient is doing well. | FSE could only duplicate the error when depressing inactive switch on coupler, no problems were found. Completed calibrations and system is operating properly. | No Injury | 20240644 | During Preparation for Use | Yes | 50 | 201
225960 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100063 | Jonathan said the laser gave an unsuccessful calibration error while trying to calibrate a 1.4 over the wire turbo elite. He tried a larger size catheter and was able to calibrate and complete the case. Patient is doing well. Jonathan Wood # 520-444-1534 | FSE was not able to duplicate the failure to calibrate, completed calibrations and system is operating properly. | No Injury | 20240652 | During Preparation for Use | Yes | 50
226022 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100180 | Cortney TM said the laser gave a Fault code 310 while using the laser. Rebooted the laser and it came up and worked properly with no fault codes. Case was completed with laser and Patient is doing well.Cortney # 301-793-7491 | FSE was not able to duplicate the reported problem, log indicated an E200 (bad catheter insertion) and no 310. Completed calibrations and system is operating properly. | No Injury | 20240651 | During Use in Patient | Yes | 49 | 310
226044 | Clinical catheter fails to calibrate | Unknown Fault | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MI | UNITED STATES (USA) | 100413 | Michelle said the laser faulted with an unable to calibrate 1.4 catheter. They tried another 1.4 catheter and it also would not calibrate, so they used a .9 and it calibrated and they were able to complete the case with the laser. Patient is doing well. *Michelle needs both of the 1.4 catheters replaced for this customer. | FSE verified failure to calibrate, completed attenuator zero and alignments. Completed calibrations and system is operating properly. | No Injury | 20240658 | During Preparation for Use | Yes | 53
226045 | The laser system indicates fault codes 400-499 - Heartbeat (HB) faults (NEX) | Fault 404 | Fault 404 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100575 | Wade Wegescheide @ 314-783-1160 said they powered up the laser and got a 404 Fault code. Case was completed with ISR Balloon and patient is doing well. | FSE was not able to duplicate the E404, replaced the HMI as a precaution. completed calibrations and system is operating properly. | No Injury | 20240659 | During Preparation for Use | Yes | 49 | 404
226116 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 231 | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100113 | TM Jamie Stowe @ 443-812-5694 said they got a Fault code 231 when they tried to calibrate their catheter, She rebooted power to the laser and was able to calibrate the catheter she seem to think it was slower to calibrate and was concerned. The case is on going now with the laser she is going to call back if there are any issues with the laser and I will update then. Jamie told me that the laser did calibrate and was ready to use but the Doctor decided to use The 13Fr TightRail to complete the case. Patient is doing well. | FSE found a loose connection to the HVPS interlock plug not completely inserted, connected. Completed calibrations and system is operating properly. | No Injury | 20240663 | During Preparation for Use | Yes | 48 | 231
226401 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 310 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100159 | E310 during the case. Error did not clear with a restart of the laser. Case was not completed with this laser. It was completed with the CVX300 without delay at the same facility. | FSE found system has low energy, completed calibrations and system is operational. The system will be deinstalled and replaced with 100631. | No Injury | 20240694 | During Use in Patient | Yes | 45 | 310
226601 | Clinical catheter fails to calibrate | Unknown Fault | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | NV | UNITED STATES (USA) | 100527 | Fail to Cal with multiple 0.9mm clinical catheters. 2.5mm reference catheter calibrates with no issue. Case was not completed with the laser. It was completed with a shockwave device without delay at the same facility. | FSE was able to verify failure to calibrate, adjusted rail alignment and replaced the attenuator. Completed calibrations and system is operating properly. | No Injury | 20240702 | During Preparation for Use | Yes | 43
226818 | Unknown fault code displayed | Electro / Mechanical | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | SC | UNITED STATES (USA) | 100192 | Rep called to report an unknown error during calibration with a 2.0 catheter, power cycled the system and error cleared. Completed procedure without further issue. | FSE found a defective CES sensor, replaced. Completed calibrations and system is operating properly. | No Injury | 20240717 | During Preparation for Use | Yes | 53
226929 | Clinical catheter fails to calibrate | Catheter Coupler Assy | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | N/A | GERMANY (DEU) | 100366 | I was Yesterday called by our customer Rhön Klinikum Bad Neustadt, that our PLS Laser System is not working, they used 2 Catheters (Turbo-Elite 2.0mm) an hat no chance to Calibrate them.On first Examination the System worked ok with the Test Catheter (Output around 55-60mj) during calibration.I guessed that we had maybe 2 faulty Catheters and took them with me for complaint.But early this morning they called me again and said that today 12th March a new catheter showed the same symptom of being unable to calibrate. So please appoint a repair of the system and replace 3 Laser catheters that where used and did not function. The 2 Laser Catheters which are property of Bad Neustadt 11MAR2024 420-006 LOT FBA21L02A292, FBA21L02A392420-006 Lot FBA23H14A 12MAR2024Update from Udo Schmid: 19MAR20241) There was a Patient in the Room2) Physician Dr. Puskas Zoltan Zoltan.puskas@campus-nes.de +49977166239503) The Procedure could not be completed with the System as it did not function, they treated the Patient in another way4) Local Anaesthesia5) Type Peripheral Treatment of the Arteria6) 2nd Patient in the Room again yes7) Local anaesthesia to8) Both patient treated on the same day treatment is unknown at the moment by me , as I can not go out to the customer at the moment due being sick. | FSE found a defective catheter coupler assembly - A34 damage, replaced. Completed calibrations and system is operating properly. | No Injury | 20240654-BV | During Preparation for Use | Yes | 38
226988 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100289 | Reported to FSE the customer experienced an E103 with a .9 catheter. Case was completed by other means during the same setting.Update from Jeff Tobin 13MAR2024During the case, when the error message first appeared, dr. Nathanson had already treated the lesion so he decided to take the laser out of the body and not do any additional atherectomy. He then ballooned and finished the case with no adverse Patient interaction. | FSE was not able to duplicate the E103, did note CES sensor was dirty and cleaned. Completed calibrations and system is operating properly. | No Injury | 20240735 | During Use in Patient | Yes | 41 | 103
227086 |  |  | Unknown | Calibration | Did not calibrate as required |  | Closed | No | NV | UNITED STATES (USA) | 100527 | Customer reported failure to calibrate with 2 - 011-004 catheters. Case was completed by other means during the same setting. |  | No Injury | N/A | During Preparation for Use |  | 0
227209 | Clinical catheter fails to calibrate | Unknown Fault | N/A | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | FL | UNITED STATES (USA) | 100559 | Rep called to report a failure to calibrate during the lead extraction, was not sure which catheter was used. Case was completed with mechanical devices during the same setting. | FSE verified failure to calibrate, found a defective PES, software issue, and damaged beam blind. Replaced all and completed calibrations and system is operating properly. | No Injury | 20240749 | During Preparation for Use | Yes | 36
227532 | The laser system has broken components | Emergency Switch | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | PA | UNITED STATES (USA) | 100053 | Ryan said he noticed the plastic cover over the Red Estop button was broke and needs to be replaced. Estop button is still functional. No case involved. | FSE verified damaged Estop, replaced. Completed calibrations and system is operating properly. | No Patient Present | 20240941 | Not during a case | Yes | 31
227734 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MS | UNITED STATES (USA) | 100600 | Philips Rep Lawrence Laurenzi @ 901-692-1815 said they got a 103 error while trying to calibrate a catheter. Case was completed by competitor - CSI. patient is doing well. | FSE was not able to duplicate, cleaned optics and completed calibrations and system is operating properly. | No Injury | 20240956 | During Preparation for Use | Yes | 29 | 103
227741 | Unknown fault code displayed | Electro / Mechanical | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | 36 | INDIA (IND) | 100095 | Customer has reported error while routine check-up during 2.5mm and 0.9mm catheter calibration. | FSE found attenuator was defective, replaced and completed calibration. system is operating properly. | No Patient Present | 20240957 | Not during a case | Yes | 29
227742 |  |  | N/A | Alarms/ Messages | Warning/ Message/ Error Message |  | Closed | No | 36 | INDIA (IND) | 100095 | Customer has reported error while routine check-up during 2.5mm and 0.9mm catheter calibration. |  | No Patient Present | 20240957 | Not during a case |  | 0
227747 | Unknown fault code displayed | Fault 221 | N/A | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100042 | Bryan Schaus TM @ 571-325-4200 said they got a System Error on their PLS before the case with no other information on the System error. They reset power to the laser and it worked properly for the case. Patient is doing well. | FSE was not able to duplicate the reported problem, did note an E221 in the system log, replaced the attenuator. Completed calibrations and system is operating properly. | No Injury | 20240960 | During Preparation for Use | Yes | 29
227775 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 300 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OK | UNITED STATES (USA) | 100159 | Philips Rep Wade Wegescheide said they got an Error code 300, 303 and 310 all at the same time just rotating between the error codes. Case was completed with back up laser and patient is doing well.This laser is in process to be replaced , just waiting for the new one to arrive. | System is slated for replacement and has not arrived yet. This system will be deinstalled. | No Injury | 20240962 | During Preparation for Use | Yes | 29 | 300
227788 | The control panel does not provide status of laser system | Wiring | N/A | Hardware | Display / Monitor issue | Non-Reportable | Closed | Yes | CA | UNITED STATES (USA) | 100024 | Karina Acorda @ 310-738-8627 Said the account powered on the laser and the display is blank. Case was completed with Balloon . Patient is doing well. Please call Karina with an eta for service. | FSE found a loose connection to the HMI, connected and completed calibrations. System is operating properly. | No Injury | 20240969 | During Preparation for Use | Yes | 32
227822 | Clinical catheter fails to calibrate | Rare Occurrence | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100366 | CS Note: this is a retroactive corrective action after the audit of the Complaint 226929. -Diana Insigne, 22 March 2024.Nature of Complaint:I was Yesterday called by our customer Rhön Klinikum Bad Neustadt, that our PLS Laser System is not working, they used 2 Catheters (Turbo-Elite 2.0mm) an hat no chance to Calibrate them.On first Examination the System worked ok with the Test Catheter (Output around 55-60mj) during calibration.I guessed that we had maybe 2 faulty Catheters and took them with me for complaint.But early this morning they called me again and said that today 12th March a new catheter showed the same symptom of being unable to calibrate. So please appoint a repair of the system and replace 3 Laser catheters that where used and did not function. The 2 Laser Catheters which are property of Bad Neustadt 11MAR2024 420-006 LOT FBA21L02A292, FBA21L02A392420-006 Lot FBA23H14A 12MAR2024Update from Udo Schmid: 19MAR20241) There was a Patient in the Room2) Physician Dr. Puskas Zoltan Zoltan.puskas@campus-nes.de +49977166239503) The Procedure could not be completed with the System as it did not function, they treated the Patient in another way4) Local Anaesthesia5) Type Peripheral Treatment of the Arteria6) 2nd Patient in the Room again yes7) Local anaesthesia to8) Both patient treated on the same day treatment is unknown at the moment by me , as I can not go out to the customer at the moment due being sick. | FSE found the system would not boot up, replaced several HMI and all PCB's, system will still not boot. System will be deinstalled and replaced. | No Injury | 20240654-BV | During Preparation for Use | Yes | 35
227890 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 240 | Fault 240 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100536 | Seth Peterson 301-367-5967 said they got an Error code 240 while trying to calibrate a catheter, rebooted system and the same error code came back when they tried to calibrate. Case was completed with balloon. Patient is doing well. Seth said they had cases scheduled for Monday of next week. | FSE was not able to duplicate the reported problem but verified in the log. Did note the HVPS connector was not clicked in all the way, seated. Completed calibrations and system is operating properly. | No Injury | 20240974 | During Preparation for Use | Yes | 33 | 240
228209 | The laser system wheels adversely affect mobility | Electro / Mechanical | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes | IL | UNITED STATES (USA) | 100327 | Customer observed broken wheel on laser. No patient involvement | FSE verified damaged caster and replaced. Completed calibrations and system is operating properly. | No Patient Present | 20241004 | Not during a case | Yes | 34
228438 | Clinical catheter fails to calibrate | Electro / Mechanical | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100414 | Laser will not calibrate using multiple catheters | FSE verified and completed atten zero routine, completed calibrations and system is operating properly. | No Injury | 20241034 | During Preparation for Use | Yes | 32
228509 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 303 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100262 | E303 with a 2.0mm catheter. Restarted the laser and calibrated a 2.3mm catheter and the laser faulted with an E106. Restarted the laser again and recalibrated the 2.3mm catheter and faulted again with an E106. Case is continuing without the laser. FSE will need to confirm with the customer how the case was completed. Sales Rep indicated that there is no injury to the patient.Update 09Apr2024: FSE confirmed with the customer that the case was completed by other means without delay at the same facility. TF 09Apr2024 | FSE was not able to duplicate the E303, did verify E103 and 303 in the system log. Completed calibrations and system is operating properly. | No Injury | 20241044 | During Use in Patient | Yes | 29 | 106
228731 | Clinical catheter fails to calibrate | Electro / Mechanical | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | VA | UNITED STATES (USA) | 100427 | Fail to cal a 0.9mm catheter. Case was not completed with this laser. The case was completed with their second laser without delay at the same facility. | FSE verified failure to calibrate, also noted E106, replaced the fiber coupler assy and also found a noisy VES, replaced. Completed calibrations and system is operating properly. | No Injury | 20241070 | During Preparation for Use | Yes | 27
228825 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | POLAND (POL) | 100505 | Fault code 221 after installation of the laser during the user training It happened with 0.9mm catheter at high fluence | During installation the FSE noted a 221, replaced the encoder motor, completed calibrations and system is operating properly. | No Patient Present | 20240708-BV | During Installation | Yes | 35 | 221
228844 | The laser system indicates fault codes 300-399 - State faults (NEX) | Fault 310 | Fault 303 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AZ | UNITED STATES (USA) | 100375 | E303/310 that would not clear. Case was not completed with the laser. The case was completed by other means without delay at the same facility. | FSE found low energy, system was deinstalled and replaced. | No Injury | 20241081 | During Preparation for Use | Yes | 35 | 303
228846 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | OR | UNITED STATES (USA) | 100288 | E500/501 during initial startup and would not clear with a restart of the laser. Case was not completed with the laser. The case was completed by other means without delay at the same facility. | FSE was not able to duplicate the E500, noted the card enclosure was backwards and corrected. Completed calibrations and system is operating properly. | No Injury | 20241082 | During Preparation for Use | Yes | 43 | 500
229082 | Multiple Fault Codes Displayed | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100262 | Error Code 106 found during the review of the System Log on 4.1.24 15:29Error Code 103 found during the review of the System Log on 4.1.24 15:29Error Code 106 found during the review of the System Log on 4.1.24 15:32Error Code 103 found during the review of the System Log on 4.1.24 15:35 | During complaint call the FSE noted a E106 and E103 in the system log, not duplicated. System is operating properly. | No Patient Present | 20241044 | Not during a case | Yes | 35 | 106
229083 | The laser system indicates fault codes 400-499 - Heartbeat (HB) faults (NEX) | Fault 401 | Fault 401 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100262 | Error Code 401 found during the review of the System Log on 4.3.24 16:40 | During complaint call the FSE noted a E401 in the system log, not duplicated. System is operating properly. | No Patient Present | 20241044 | Not during a case | Yes | 39 | 401
229112 | The laser system generates excessive noise | E-Stop Engaged During Power-Up | Unknown | Alarms/ Messages | Constant alarm | Non-Reportable | Closed | Yes |  | UNITED KINGDOM (GBR) | 100297 | When the customer powered on the system for a procedure, they had an audible alarm sounding. They tried to clear the alarm by resetting the system CB1 circuit breaker.After the customer could not get the alarm to clear they cancelled the procedure.When they contacted me I ran through some troubleshooting and the emergency off button was depressed. Once this had been reset the system was operating as expected. No further action requiredADDITIONAL INFORMATION OBTAINED 09APR2024:• Since the case was aborted, did the patient have to be transported to another facility to receive treatment? No• Did the facility and the physician have alternate options to treat the patient, in case the patient needed them? Yes• Did aborting the case place the patient at an increased risk? No• Was the patient open on the table and access obtain? Yes (open is not an appropriate choice of words for a percutaneous procedure but radial access was obtained).• Was the patient already under anesthesia or sedation? Local anaesthetic only• Was angiogram performed? Yes• Was this a LM, Coronary, or peripheral case? CoronaryADDITIONAL INFORMATION OBTAINED 11APR2024:What happened to/or what was the plan for the patient after the case was aborted? Balloon angioplasty was performed with the plan to return at later date for Laser procedure What was the indication for the procedure? Repeated in stent restenosis to Prox/Ostial RCA. | TS determined E stop was engaged during power up, system is operating properly. | No Injury | N/A | Preparation for use, after initial prep | No | 36
229244 | The laser system wheels adversely affect mobility | Compromised Wheels | N/A | Hardware | Component or Part Issue | Non-Reportable | Closed | Yes |  | ITALY (ITA) | 100008 | Sales Manager of Newtech contacted our FSE.Posterior wheel of PLS 100008 is broken.Image shared by FSEUpdate: Hospital put the part back into place and no problem shows. Nonetheless FSE will go to the hospital and replace the wheel. | FSE verified damaged caster and replaced. completed calibrations and system is operating properly. | No Patient Present | N/A | Not during a case | Yes | 36
229246 | Clinical catheter fails to calibrate |  | Unknown | Calibration | Did not calibrate as required | Non-Reportable | Open | Yes |  | SPAIN (ESP) | 100004 | They had 3 laser lead extraction cases planned.Initially they calibrated correctly with the reference catheter but the calibration of the glidelight (14F) failed.They turned the laser equipment on and off several times and calibrated several more times, they even opened another glidelight catheter (in case the error came from the catheter) but it also failed, and after several times of turning it on and off the calibration was ok.For the second and third patients the problem was the same. The reference catheter calibrated ok but the sterile catheter failed, and after several on, off and recalibrations with the sterile catheter the calibration was fine. Something similar has happened in peripheral vascular surgery, which was reported to us yesterday afternoon.Distributor contacted to provide forms for all cases.The problem was last Friday. They did correct calibration of the reference catheter, in this case the calibration with the turboelite was also ok, and in the middle of the procedure the system failed. They turned it off, turned it on, recalibrated and it worked again.CS note: Complaint form received today 24 April 2024, please see attached. | FSE found an E221 during checks, replaced the encoder and completed calibrations. System is operating properly. | No Injury | 20240703-BV | During Use in Patient | Yes | 42
229341 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | AL | UNITED STATES (USA) | 100369 | Rep reported an E106 with a .9 catheter, completed the case using a 1.4 without issue. | FSE was not able to duplicate the reported problem, completed calibrations and system is operating properly. | No Injury | 20241102 | During Use in Patient | Yes | 34 | 106
229437 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100590 | Per Rob, clinical went to turn on PLS and immediately gave a 221. Was able to cycle power to PLS and clear error. Case was completed without further issue. | FSE found a defective attenuator, replaced. Completed calibrations and system is operating properly. | No Injury | 20241109 | During Preparation for Use | Yes | 33 | 221
229564 | Clinical catheter fails to calibrate | Electro / Mechanical | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MI | UNITED STATES (USA) | 100353 | Chris Dahlem-(Clinical rep) Said they were getting an Unable to Calibrate with the 1.4 clinical catheter. He rebooted and tried again, and still got the Unable to Calibrate error. He said this laser just had the software upgrade last friday. Case was completed with a competitors laser (Reone Laser). Patient is doing well. | FSE completed attenuator zero and calibrations and system is operating properly. | No Injury | 20241265 | During Preparation for Use | Yes | 30
229664 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100590 | Carolynn was doing an in-service on this laser and she got a 106 error code. No case , no patient involved. | FSE was not able to duplicate the reported problem and replaced the VES as a precaution. Completed calibrations and system is operating proeprly. | No Patient Present | 202411272 | Not during a case | Yes | 30 | 106
229755 | The reference catheter fails to calibrate | Fault 108 | Fault 108 | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes |  | SPAIN (ESP) | 100021 | A compliant has been opened for an instance of error 108 which happened two days before service of the laser. BIOMENCO notified FSE same day the issue happened and as he was servicing a laser in Valencia of the same distributor's network. log Analysis shows that error 108 is preceded by Error 321 "NOT_IN_READY_STATE_PROCEDURE" which means that the operator didn't press "SET TO READY" button prior to depressing the foot switch. No action taken other than SW Update and routine PM as scheduled since the error 108 is supposed to be corrected by performing the SW upgrade. Preventive Maintenence and electrical safety tests performed. Completed FCO CR175109 Software upgrade v 1.0.1.0. The system matches Philips/Spectranetics specifications. PLS System is ok for use.While performing maintenance on the laser system encountered that the Hospital's reference catheter of 2.5 mm was malfunctioning. After FSE used his own 2.5 mm catheter, maintenance could continue.UPDATE: Laser fault occured while in use at the end of the procedure, no injury or issues occured. Balloon was used.Sent an order to deliver a new catheter to distributor Biomenco.- Biomenco mentioned there was a patient on the table but no furhter information was provided by them. Asked to contact Karen from Biomenco for further information. | FSE updated the software and completed calibrations, system is operating properly. | No Injury | 20230652-BV | During Use in Patient | Yes | 29 | 108
229774 | Clinical catheter fails to calibrate | Electro / Mechanical | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | MA | UNITED STATES (USA) | 100410 | Nicole McCarthy @ 401-464-1725 said they got a No Energy/ System failure error while trying to calibrate .9 clinical catheter, they tried the Ref catheter and it calibrated successfully, They also tried a 2.0 clinical catheter after that and it also gave them a No Energy/System Failure. Case was completed with Balloon. patient is doing well. | FSE replaced the attenuator motor, completed calibrations and system is operating properly. | No Injury | 20241278 | During Preparation for Use | Yes | 30
229868 | Multiple Fault Codes Displayed | Fault 106 | Fault 302 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | GERMANY (DEU) | 100082 | Today the FC 302. 106 happened on the PLS 100082 with a 2.3 TurbopowerThe fault happens again when changing the catheter.==============================================================================FSE is dispatched asap to assess the situation and try solving the problem on site.FSE Will provide a customer complaint form to the customer to increase data collected for the complaint in CATSWEB.==============================================================================During the treatment the laser was failing with Fault 106 and 302. Immediate service needed. 2x Sterile Cath opened | FSE stated the attenuator was off by 2 steps, completed attenuator zero and completed calibrations. System is operating properly. | No Injury | 20240747-BV | During Use in Patient | Yes | 29 | 302
229870 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Electro / Mechanical | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | NJ | UNITED STATES (USA) | 100549 | Katya Oretga 201-602-1931 said they are getting a 103 error code when the laser is set to 60/60, if they go below 60/60 the laser works properly. They tried a second catheter and was able to complete the case but it would also error 103 at 60/60. Patient is doing well. | FSE found debris on the CES detector, cleaned. Completed calibrations and system is operating properly. | No Injury | 20241280 | During Use in Patient | Yes | 29 | 103
230011 | The laser system does not provide reliable performance | Fault 310 | N/A | Other | Other | Non-Reportable | Closed | Yes | NY | UNITED STATES (USA) | 100128 | Kyle Riccardi - Recently in a lead extraction, we experienced the error where the system freezes and has to be turned off and then on again after the physician accidentally steps on the pedal too quickly between 10 seconds of laser time/during its 5 second buffer time. Case was completed with laser and is doing well. | FSE found the footswitch causing intermittent error, replaced. Competed calibrations and system is operating proprely. | No Injury | 20241288 | During Use in Patient | Yes | 28
230053 |  |  | Fault 500 |  |  |  | Closed | No |  | SWITZERLAND (CHE) | 100309 | Claudia del Fatti called FSE, Olaf Bruns for PLS in Zurich which shwoed FC 500. Critical patient on the table and hospitl ersonnel tried rebooting the laser a few times with no solution.FC 500 indicates an initialisation failure, which could mean 1 of 3 detectors is failing, the attenuator motor or the shutter.Expect to receive complaint form from Claudia |  | No Injury | N/A | Unknown |  | 1 | 500
230054 | The laser system indicates fault codes 500-599 - Software faults (NEX) | Fault 500 | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes |  | SWITZERLAND (CHE) | 100309 | Claudia del Fatti called FSE, Olaf Bruns for PLS in Zurich which shwoed FC 500. Critical patient on the table and hospitl ersonnel tried rebooting the laser a few times with no solution.FC 500 indicates an initialisation failure, which could mean 1 of 3 detectors is failing, the attenuator motor or the shutter.Expect to receive complaint form from ClaudiaADDITIONAL INFORMATION OBTAINED 23APR2024:The customer will not report the case.They did treatment with this cohort of patients before they had the laser available, and they could do it just the usual standard way they did treatment before.No extras except they needed to change to other tools use strategy instead of using the laser.Patient went out of the OR well, no delayed outcome.Patient released the same day | FSE verified and found the PES sensor was defective, replaced. Completed calibrations and system is operating properly. | No Injury | 20240750-BV | During Preparation for Use | Yes | 31 | 500
230138 | Clinical catheter fails to calibrate | Electro / Mechanical | Unknown | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | Az | UNITED STATES (USA) | 100358 | During the case a .9 catheter failed to calibrate. Two separate .9 catheters were tried. Staff then went to a 1.4 catheter and it calibrated. Case was completed successfully with 1.4 catheter. | FSE was not able to duplicate and replaced the attenuator as a precaution. Completed calibrations and system is operating properly. | No Injury | 20241299 | During Preparation for Use | Yes | 25
230289 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes | TX | UNITED STATES (USA) | 100641 | E106 when trying to use the laser | FSE found the attenuator step required adjustment, completed atten zero routine and calibrations and system is operating properly. | No Injury | 20241309 | During Preparation for Use | Yes | 29 | 106
230529 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) | Fault 106 | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Closed | Yes | GA | UNITED STATES (USA) | 100199 | Dr made two passes and both times an E106 occurred. Dr determined passes were a success and case was complete. | FSE replaced the attenuator and completed calibrations, system is operating properly. | No Injury | 20241324 | During Use in Patient | Yes | 22 | 106
230632 | The laser system generates unusual noise | E-Stop Engaged During Power-Up | Unknown | Hardware | Component or Part Issue | Non-Reportable | Open | Yes | FL | UNITED STATES (USA) | 100045 | Laser Making funny noise and monitor is blank | FSE found the e-stop was engaged during power up, released and system powered up and operated properly. | No Injury | 20241328 | During Preparation for Use | Yes | 26
230812 | Clinical catheter fails to calibrate | Debris On Focus Lens | N/A | Calibration | Unable to Calibrate | Non-Reportable | Closed | Yes | TX | UNITED STATES (USA) | 100614 | Laser does not recognize any catheters. No calibration of catheters was possible. Case was not completed with the laser. The case was completed with their CVX300 without delay at the same facility. | FSE determined the problem was failure to calibrate, found debris on focus lens, cleaned and completed calibrations. System is operating properly. | No Injury | 20241348 | During Preparation for Use | Yes | 20
231018 | Clinical catheter fails to calibrate | Electro / Mechanical | N/A | Calibration | Unable to Calibrate | Non-Reportable | Open | Yes | TX | UNITED STATES (USA) | 100081 | Fail to calibrate a 12fr & 14fr catheter. Laser power was reset and catheters calibrated. Case was completed with the laser without delay at the same facility. | FSE was not able to duplicate the failure to calibrate. Completed attenuator zero routine and calibrations, system is operating properly. | No Injury | 20241356 | During Preparation for Use | Yes | 20
231093 | The reference catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Did not calibrate as required | Non-Reportable | Closed | Yes | MD | UNITED STATES (USA) | 100247 | Failure to calibrate the 2.5mm reference catheter. Case was not completed with this laser. The case was completed with their CVX300 without delay at the same facility. | FSE was not able to duplicate the reported problem, adjusted the attenuator motor and completed calibrations. System is operating properly. | No Injury | 20241363 | During Preparation for Use | Yes | 19
231118 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) |  | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes | OR | UNITED STATES (USA) | 100116 | E231 after a few seconds of lasing. Laser was reset with the same result. Case was not completed with the laser. The case was completed with a shockwave product without delay at the same facility. |  | No Injury | 20241365 | During Use in Patient | Yes | 19 | 231
231163 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) |  | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes | FL | UNITED STATES (USA) | 100312 | Rep called to report an error 103 after calibration, customer rebooted the system and completed the case without issue. |  | No Injury | 20241366 | During Use in Patient | Yes | 16 | 103
231181 | Clinical catheter fails to calibrate |  | N/A | Calibration | Unable to Calibrate | Non-Reportable | Open | Yes | MA | UNITED STATES (USA) | 100410 | Rep called to report a failure to calibrate with 2 different .9mm catheters. Reference catheter calibrated and a 1.4mm catheter calibrated which was used to complete the case. |  | No Injury | 20241370 | During Preparation for Use | Yes | 16
231218 | The laser system indicates fault codes 300-399 - State faults (NEX) |  | Fault 321 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes |  | GERMANY (DEU) | 100001 | Laser system is down with a patient on the table. The patient was safe and the case was succesful using a tightrail. |  | No Injury | 20240768-BV | During Use in Patient | Yes | 15 | 321
231285 | Clinical catheter fails to calibrate | Unable to replicate complaint | N/A | Calibration | Unable to Calibrate | Non-Reportable | Open | Yes | PA | UNITED STATES (USA) | 100100 | Rep called to report a failure to calibrate, was not sure which catheter the customer was using. Case was completed by other means during the same setting. | FSE was not able to duplicate but noted a bad catheter, completed calibrations and system is operating properly. | No Injury |  | During Preparation for Use | Yes | 15
231399 | Clinical catheter fails to calibrate |  | N/A | Calibration | Unable to Calibrate | Non-Reportable | Open | Yes | AK | UNITED STATES (USA) | 100530 | Rep called to report a failure to calibrate with a .9, 1.7, and a 2.0. Calibrates fine with the 2.5 reference catheter. Case completed by other means during the same setting. |  | No Injury |  | During Preparation for Use | Yes | 14
231510 | Clinical catheter fails to calibrate |  | N/A | Calibration | Unable to Calibrate | Non-Reportable | Open | Yes | NJ | UNITED STATES (USA) | 100264 | Customer called to report a failure to calibrate with several catheters during testing prior to cases today. No patient involved in this event. The customer has a full day of cases today. | FSE was not able to duplicate the reported problem, completed calibrations and system is operating properly. | No Patient Present |  | Not during a case | Yes | 13
231669 | Clinical catheter fails to calibrate |  | N/A | Calibration | Unable to Calibrate | Non-Reportable | Open | Yes | FL | UNITED STATES (USA) | 100421 | Can I request a PLS complaint is created linked to Turbo Elite Complaint 228698?We received the Turbo Elite device back with reported complaint: Catheter worked great first time however during the second time the catheter would not recalibrate. A new turbo Elite was used to complete the case.During the Turbo Elite evaluation, it was observed that the foot clamp within the proximal coupler was damaged (see below photo). It appears that the laser burned a hole through the foot clamp, and the fibers were visible beneath. The foot clamp was sent to the Failure Analysis Lab for testing to determine if the clamp material was per spec. The lab determined that the metal was a close match for Stainless Steel 301 and 304, so we requested CofCs to confirm. We received the CofCs today, 08MAY2024. The CofCs for foot clamp 3650-1220 indicate the metal is Stainless Steel T301, which is in alignment with failure analysis lab results. As a result, we believe the damage observed is caused by the PLS. |  | No Injury |  | After the Procedure | Yes | 12
231684 | The laser system indicates fault codes 500-599 - Software faults (NEX) |  | Fault 500 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes | OR | UNITED STATES (USA) | 100370 | Customer reported an E500, case was completed by other means during the same setting. |  | No Injury | N/A | During Preparation for Use | Yes | 12 | 500
231737 | The laser system has malfunctioning components | Electro / Mechanical | N/A | Hardware | Component or Part Issue | Non-Reportable | Open | Yes |  | JAPAN (JPN) | 100569 | FSE performed incoming inspection for PLS #100569 and found that S1 switch didn't work. He then removed the coupler and found that the S1 switch on the A34 board was damaged.Was the Device in Clinical Use at the Time the Issue was discovered? -> NoWas any Patient or User Harmed? -> No | During installation the FSE noted the A34 S1 was damaged, replaced the coupler. Completed calibrations and system is operating properly. | No Patient Present | 20241395 | Not during a case | Yes | 9
231744 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) |  | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes |  | SPAIN (ESP) | 100004 | This complaint is being filed with retroactive action since the hospital did not warn us or the distributor plus when asked the distributor to file it, they metioned that the hospital might not provide us with the correct data since many teams use the PLS and it happened a long time ago.During the service, when download and analyze the LOG Files of the PLS it was indicating FC 221 on the dates: 14. FEB 24, 28. FEB 24, 18. MAR. 24, 19.MAR. 24. This faults hadn't been reported to us or the distributor. I was told that they rebooted the PLS and performed the operations properly. We didn'T got back a feedback about the kind of the case or datas from the customer. Our distributor Biomenco tried to get some information, but this was declined by the customer, but there was no incident with the patient | During complaint call the FSE noted a E221 in the system log, not duplicated, completed calibrations and system is operating properly. | No Patient Present | 20240703-BV | Not during a case | Yes | 9 | 221
231745 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) |  | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes |  | SPAIN (ESP) | 100004 | This complaint is being filed with a retroactive action since the hospital did not warn us on the distributor plus when asked the distributor to file it, they mentioned that the Hospital might not provide us with the correct data since many teams use the PLS and it happened a long time ago.During the service, when download and analyze the LOG Files of the PLS it was indicating FC 221 on the dates: 14. FEB 24, 28. FEB 24, 18. MAR. 24, 19.MAR. 24. This faults hadn't been reported to us or the distributor. I was told that they rebooted the PLS and performed the operations properly. We didn'T got back a feedback about the kind of the case or datas from the customer. Our distributor Biomenco tried to get some information, but this was declined by the customer, but there was no incident with the patient | During complaint call the FSE noted a E221 in the system log, not duplicated, completed calibrations and system is operating properly. | No Patient Present | 20240703-BV | Not during a case | Yes | 9 | 221
231750 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) |  | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes |  | SPAIN (ESP) | 100004 | This complaint is filed with retroactive action since the hospital did not warn us or the distributror plus when asked to file it, the distributor mentioned that the hospital would not provide us with the correct data since many teams use the PLS plus the issue happened a long time ago.During the service, when download and analyze the LOG Files of the PLS it was indicating FC 221 on the dates: 14. FEB 24, 28. FEB 24, 18. MAR. 24, 19.MAR. 24. This faults hadn't been reported to us or the distributor. I was told that they rebooted the PLS and performed the operations properly. We didn'T got back a feedback about the kind of the case or datas from the customer. Our distributor Biomenco tried to get some information, but this was declined by the customer, but there was no incident with the patient | During complaint call the FSE noted a E221 in the system log, not duplicated, completed calibrations and system is operating properly. | No Patient Present | 20240703-BV | Not during a case | Yes | 9 | 221
231758 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) | Fault 221 | Fault 221 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes |  | SPAIN (ESP) | 100004 | This complaint is being filed with retroactive action since the hospital did not warn us or the distributor plus when asked, the distributor mentioned that the hospital might not provide us with the correct data since many different teams use it as well as the issue occured a long time ago.During the service, when download and analyze the LOG Files of the PLS it was indicating FC 221 on the dates: 14. FEB 24, 28. FEB 24, 18. MAR. 24, 19.MAR. 24. This faults hadn't been reported to us or the distributor. I was told that they rebooted the PLS and performed the operations properly. We didn'T got back a feedback about the kind of the case or datas from the customer. Our distributor Biomenco tried to get some information, but this was declined by the customer, but there was no incident with the patient | During complaint call the FSE noted a E221 in the system log, not duplicated, completed calibrations and system is operating properly. | No Patient Present | 20240703-BV | Not during a case | Yes | 9 | 221
231759 |  |  | Fault 221 |  |  |  | Open | No |  | SPAIN (ESP) | 100004 | This complaint is being filed with retroactive action since the hospital did not warn us or the distributor plus when asked, the distributor mentioned that the hospital might not provide us with the correct data since many different teams use it as well as the issue occured a long time ago.During the service, when download and analyze the LOG Files of the PLS it was indicating FC 221 on the dates: 14. FEB 24, 28. FEB 24, 18. MAR. 24, 19.MAR. 24. This faults hadn't been reported to us or the distributor. I was told that they rebooted the PLS and performed the operations properly. We didn'T got back a feedback about the kind of the case or datas from the customer. Our distributor Biomenco tried to get some information, but this was declined by the customer, but there was no incident with the patient |  | No Patient Present | 20240703-BV | Not during a case |  | 9 | 221
231812 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) |  | Fault 106 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes | NC | UNITED STATES (USA) | 100125 | E106 when trying to calibrate |  | No Injury |  | During Preparation for Use | Yes | 9 | 106
231928 | Clinical catheter fails to calibrate |  | N/A | Calibration | Unable to Calibrate | Non-Reportable | Open | Yes | NJ | UNITED STATES (USA) | 100264 | Erin said there laser would not calibrate , Case was completed with their CVX laser. patient is doing well. |  | No Injury | 20241417 | During Preparation for Use | Yes | 8
231967 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) |  | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes | FL | UNITED STATES (USA) | 100578 | Chris said they got a 103 error code when they tried to calibrate the clinical catheter. They reset laser and re inserted the catheter and was able to calibrate and start the case , they completed 6000 pulses and the laser gave them the 103 error code again. they completed the case with Balloon. Patient is doing well. |  | No Injury | 20241422 | During Use in Patient | Yes | 8 | 103
231974 | The laser system indicates fault codes 600-699 - Temperature faults (NEX) |  | Fault 600 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes | OK | UNITED STATES (USA) | 100372 | Mike said they got a 600 error code , They used their other laser to complete the case. Patient is doing well. |  | No Injury | 20241556 | During Preparation for Use | Yes | 8 | 600
232070 | The laser system indicates fault codes 100-199 - Calibration faults (NEX) |  | Fault 103 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes | FL | UNITED STATES (USA) | 100421 | ROb said they got a 103 error code with 2 1.4 clinical catheters and he also tried a 1.7 clinical catheter as was able to get past the 103 error code for the calibration but it came came back as they started to use it for the case. Case was completed with Balloon.Patient is doing well.Please call Rob with eta for service as they have 6 cases scheduled for tomorrow |  | No Injury | 20241560 | During Preparation for Use | Yes | 7 | 103
232563 |  |  | Fault 231 | Alarms/ Messages | Warning/ Message/ Error Message |  | Open | Yes | CO | UNITED STATES (USA) | 100252 | error 230 and 231 during a case today. The rep restarted it and was able to calibrate the ref cath. They had to stop the case for a reason unrelated to the laser.ADDITIONAL INFORMATION OBTAINED 22MAY2024:Yes, patient was under anesthesia. The case was stopped because we had made an attempt to remove the lead but could not with our also removing another lead on a dependent patient. So, the physician decided to stop extracting and just put in one new lead.Yes, the laser did recalibrate. I turned the PLS off and back on. Calibrated the reference catheter which did work. Then calibrated the 16Fr GlideLight laser sheath which also worked. The error came up during use of the 16Fr GL. This was after using a 12Fr GL & 14Fr GL earlier in the case. |  | No Injury |  | During Preparation for Use |  | 1 | 231
232570 |  |  | Fault 230 |  |  |  | Open |  |  | POLAND (POL) | 100505 | Fault code 230 and 103 during the user training It happened with 0.9 mm catheter at high fluence 60 mJ/mm2 |  | No Patient Present | 20240824-BV | Not during a case |  | 1 | 230
232604 | The laser system indicates fault codes 200-299 - Hardware faults (NEX) |  | Fault 200 | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes |  | JAPAN (JPN) | 100500 | The customer reported that error 200 occurred during preparation for use.Was the Device in Clinical Use at the Time the Issue was discovered? -> NoWas any Patient or User Harmed? -> NoFES will visit the site and investigate the case. |  | No Injury |  | Not during a case | Yes | 0 | 200
232650 | Unknown fault code displayed |  | Unknown | Alarms/ Messages | Warning/ Message/ Error Message | Non-Reportable | Open | Yes |  | INDIA (IND) | 100302 | During routine service; FCO 1010 of system. |  | No Injury | 20240997 | Not during a case | Yes | 0